Table Of Contents

3.2.P.3.5 process Validation and/or Evaluation	7

1.	Summary	7

2.	Drug Product Solution Filtration	7

3.	Component Depyrogenation	10

3.1	Depyrogenation Validation Study Summary	10

3.2	Validation and Production Depyrogenation Parameters	11

3.3	Initial Qualification	11

3.3.1	Load Pattern, Probes and Endotoxin Indicator Location Details	13

3.3.2	Validation Study Data	14

3.4	Requalification	17

3.4.1	Load Pattern, Probes and Endotoxin Indicator Location Details	17

3.4.1	Validation Study Data	19

3.5	Acceptance Criteria	20

3.6	Conclusion	21

4	Component Sterilization	21

4.1	Component Sterilization Validation Study Summary	21

4.2	Validation and Production Component Sterilization Parameters	23

4.3	Initial Qualification	23

4.3.1	Load Pattern, Probes and Biological Indicator Location Details	24

4.3.2	Validation Study Data	25

4.4	Requalification	29

4.4.1	Load Pattern, Probes and Biological Indicator Location Details	29

4.4.2	Validation Study Data	30

4.5	Acceptance Criteria	33

4.6	Conclusion	34

5	Component Sterilization (RTU)	34

5.1	Seal Sterilization Validation Study Summary	35

5.1.1	Dose Mapping Study	35

5.1.2	Dosimetry Study	37

6	EQUIPMENT Sterilization (Autoclave)	39

6.1	AUTOCLAVE (XX-XXX)	39

6.1.1	Equipment Sterilization Validation Study Summary	39

6.1.2	Validation and Production Equipment Sterilization Parameters	41

6.1.3	Initial Qualification	42

6.1.3.1	Load Pattern, Probes and Biological Indicator Location Details	42

6.1.3.2	Validation Study Data	44

6.1.4	Requalification	47

6.1.4.1	Load Pattern, Probes and Biological Indicator Location Details	47

6.1.4.2	Validation Study Data	48

6.2	Acceptance Criteria (XX-XXX)	50

6.3	Conclusion	51

7	Filtration vessel Sterilization (Sterilization in Place)	51

7.1	Equipment Sterilization in Place Validation Study Summary	51

7.2	Validation and Production Equipment Sterilization Parameters	52

7.3	Initial Qualification	53

7.3.1	Probes and Biological Indicator Location Details	53

7.3.2	Validation Study Data	54

7.4	Requalification	55

7.4.1	Probes and Biological Indicator Location Details	55

7.4.2	Validation Study Data	55

7.5	Acceptance Criteria	56

7.6	Conclusion	57

8	LYOPHILIZer STERILIZATION (Sterilization in place)	57

8.1	Lyophilizer	57

8.1.1	Equipment Sterilization in Place Validation Study Summary	57

8.1.2	Validation and Production Equipment Sterilization Parameters	58

8.1.3	Initial Qualification	58

8.1.3.1	Probes and Biological Indicator Location Details	59

8.1.3.2	Validation Study Data	59

8.1.4	Requalification	60

8.1.4.1	Probes and Biological Indicator Location Details	61

8.1.4.2	Validation Study Data	61

9	Media Fill (Aseptic Process Simulation)	62

9.1	Introduction	62

9.2	Media Fill Validation Study Summary	63

9.3	Validation and Exhibit and Commercial Equipment/Component Summary	64

9.4	Validation and Exhibit and Commercial Filling Summary	66

9.5	Initial Qualification	68

9.5.1	Filled Unit Reconciliation Summary	68

9.6	Requalification	69

9.6.1	Filled Unit Reconciliation Summary	69

9.7	Acceptance Criteria	70

9.8	Conclusion	71

9.9	Media Fill Failure Investigation and Corrective Action	71

10	Terminal Sterilization	72

10.1	Terminal Sterilization Validation Study Summary	72

10.2	Validation and Production Terminal Sterilization Parameters	73

10.3	Initial Qualification	74

10.3.1	Load Pattern, Probes and Biological Indicator Location Details	75

10.3.2	Validation Study Data	75

10.4	Requalification	78

10.5	Acceptance Criteria	78

10.6	Conclusion	78

11	Container Closure Integrity validation	80

10.1	Introduction	80

10.2	Conclusion	81

10.3	Methods and controls to monitor production cycles	81

12	Exhibit Batch Filling Data	82

13	Holding periods	83

13.1	Conclusion	85

List of In-Text Tables

Table 1	Summary of Filter Integrity Test Results for Product Name	9

Table 2	Summary of Filter Validation Studies	9

Table 3	Summary: Depyrogenation Process Validation Studies	10

Table 4	Depyrogenation Parameters: Validation Vs. Production for XX ml_XX mm Clear Vial	11

Table 5	Vial specification: Validation Vs. Production for XX ml_XX mm Clear Vial		12

Table 6	Empty Chamber Heat Distribution Study Data	14

Table 7	Heat Penetration/Endotoxin Challenge Study Data: XX ml_XX mm Clear Vial		15

Table 8	Endotoxin Challenge Data	17

Table 9	Heat Penetration/Endotoxin Challenge Study Data: XX ml_XX mm Clear Vial		19

Table 10	Endotoxin Challenge Data	20

Table 11	Summary: Component Sterilization Validation Studies <As example 20 mm rubber stopper & 20 mm seals are mentioned>	22

Table 12	Autoclave Cycle Parameters: Validation Vs. Production	23

Table 13	Empty Chamber Heat Distribution Study Data	25

Table 14	20 mm Rubber Stopper and seals (Maximum Load # XXX): Heat Distribution/Heat Penetration Data	26

Table 15	20 mm Rubber Stopper and seals (Minimum Load # XXX): Heat Distribution/Heat Penetration Data	27

Table 16	20 mm Rubber Stopper and seals for liquid Product: Microbiological Efficacy by Biological Indicator Study  Data	28

Table 17	Empty Chamber Heat Distribution Study Data	30

Table 18	XX mm Rubber stopper and Aluminum seals (Maximum Load – XXX)]: Heat Distribution/Heat Penetration Data	31

Table 19	XX mm Rubber stopper and Aluminum seals (Maximum Load – XXX)]: Microbiological Efficacy by Biological Indicator Study Data	32

Table 20	Summary: Seal Sterilization Validation Studies.	35

Table 21	Dose Distribution in the shelves	38

Table 22	Summary: Equipment Sterilization Validation Studies (Autoclave XX-XXX)		40

Table 23	Autoclave Cycle Parameters: Validation Vs. Production (Autoclave XX-XXX)		41

Table 24	Empty Chamber Heat Distribution Study Data (Autoclave XX-XXX)	44

Table 25	Filling and Capping Assembly (Maximum Load): Heat Distribution/Heat Penetration Studies Data (Autoclave XX-XXX)	45

Table 26	Filling and Capping Assembly (Minimum Load): Heat Distribution/Heat Penetration Studies Data (Autoclave XX-XXX)	46

Table 27	Biological Indicator Study Summary Data: Filling and Capping Assembly Load (Autoclave XX-XXX)	47

Table 28	Filling and Capping Assembly (Maximum Load): Heat Distribution/Heat Penetration Studies Data (Autoclave XX-XXX)	48

Table 29	Biological Indicator Study Summary Data: Filling and Capping Assembly Load (Autoclave XX-XXX)	49

Table 30	Summary: Equipment Sterilization in Place Validation Studies	51

Table 31	Sterilization in Place Cycle Parameters: Validation Vs. Production	52

Table 32	SIP Validation-Temperature Profile Data: XXX L Vial filtration vessel	54

Table 33	Biological Indicators Challenge Data: XXX L Vial filtration vessel	54

Table 34	SIP Validation-Temperature Profile Data: XXX L Vial filtration vessel	55

Table 35	Biological Indicators Challenge Data: XXX L Vial filtration vessel	56

Table 36	Summary: Equipment Sterilization in Place Validation Studies	57

Table 37	Sterilization in Place Cycle Parameters: Validation Vs. Production	58

Table 38	SIP Validation-Temperature Profile Data: Lyophilizer	59

Table 39	Biological Indicators Challenge Data: Lyophilizer	60

Table 40	SIP Validation-Temperature Profile Data: Lyophilizer	61

Table 41	Biological Indicators Challenge Data: Lyophilizer	62

Table 42	Summary: Media Fill Validation Studies	64

Table 43	Equipment & Component Details: Validation Vs. Exhibit and Commercial (XX mL Vial)	64

Table 44	Comparative specification of physical dimensions for the vials	65

Table 45	Filling Summary: Validation Vs. Exhibit and Commercial: XX mL Vials (Initial Qualification)	66

Table 46	Filling Summary: Validation Vs. Exhibit and Commercial: XX mL Vials (Requalification)	67

Table 47	Filled Unit Reconciliation: X mL	68

Table 48	Filled Unit Reconciliation: X mL	69

Table 49	Summary: Terminal Sterilization Process Validation Studies	73

Report # XXX	73

Table 50	Terminal Sterilization Parameters: Validation Vs. Production	74

Table 51	Empty Chamber Heat Distribution Study Data (Load # XX)	76

Table 52	XX mL PP Vial (Maximum Load# XX): Heat Penetration Study with BI challenge run	76

Table 53	XX mL Vial (Minimum Load# XX): Heat Penetration Study with BI challenge run	77

Table 54	Microbiological Efficacy by Biological Indicator Study Data	77

Table 55	Container closure system for Product Name	80

Table 56	Acceptance Criteria and observation: Dye Ingress-ANDA Submission Batch of Product Name	81

Table 57	Exhibit Batch filling results of Product Name (Filling stage: Initial, Middle & End)		82

Table 58	Details of Holding Time Study	83

Table 59	Exhibit Batch bulk hold time results of Product Name (Unfiltered bulk (store at 60°C): Bulk solution 0 hour, 12 hours, 24 hours & 36 hours hold from end of compounding)	84

Table 60	Exhibit Batch bulk hold time results of Product Name (Filtered bulk (Store at 60°C): Bulk solution 0 hour, 12 hours, 24 hours & 36 hours hold)	84

Table 61	Exhibit Batch bulk hold time results of Product Name (Filtered bulk (Store at NMT 25°C): Bulk solution 0 hour, 12 hours, 24 hours & 36 hours hold)	84

List of In-Text Figures

Figure 1.	Probe Location for Heat Distribution Run	13

Figure 2.	Probe & Endotoxin Location for Heat Penetration and Endotoxin Challenge Study: XX ml_XX mm Clear Tubular Vial	13

Figure 3.	Probe & Endotoxin Location for Heat Penetration and Endotoxin Challenge Study: XX ml_XX mm Clear Tubular Vial	18

Figure 4.	Probe Location for Empty chamber Heat distribution Run	24

Figure 5.	20 mm Stopper and seals (Maximum load # XXX): Load Arrangement, Probe/Biological Indicator Location for Heat Distribution/Penetration Runs	24

Figure 6.	20 mm Stopper and seals (Minimum load # XXX): Load Arrangement, Probe/Biological Indicator Location for Heat Distribution/Penetration Runs	25

Figure 7.	Probe Location for Empty chamber Heat distribution Run	30

Figure 8.	 XX mm Rubber stopper and Aluminum seals (Maximum Load – XX)]: Load Arrangement, Probe/Biological Indicator Location for Heat Distribution/Penetration Runs	30

Figure 9.	Placement of Dosimeters in tote box	36

Figure 10.	Box Loading Pattern	36

Figure 11.	Placement of Dosimeters in X set	38

Figure 12.	Probe Location for Empty chamber Heat Distribution Run	42

Figure 13.	Filling and Capping Assembly Load (Maximum Load): Load Arrangement for Heat Distribution/Penetration Runs	43

Figure 14.	Filling and Capping Assembly Load (Minimum Load): Load Arrangement for Heat Distribution/Penetration Runs	43

Figure 15.	Filling and Capping Assembly Load (Maximum Load): Load Arrangement for Heat Distribution/Penetration Runs	48

Figure 16.	Diagrammatic representation of Probe/BI location for XXX L Vial filtration vessel	53

Figure 17.	Diagrammatic representation of Probe/BI location for XXX L vessel	55

Figure 18.	Diagrammatic representation of Probe/BI location for Lyophilizer	59

Figure 19.	Diagrammatic representation of Probe/BI location for Lyophilizer	61

Figure 20.	Probe Location for Empty chamber Heat Distribution Run	75

# 3.2.P.3.5 process Validation and/or Evaluation

# Summary

[Product Name] is [Add Description]. The proposed drug product is manufactured by “[add method of sterilization]” approach on the vial filling line located at [XXX] floor of [Company name].

# Drug Product Solution Filtration

<Add product specific parameters>

- The bulk drug product is filtered using [Capsule Filter Type/Size] (Catalog No. [Filter Catalog No.]) (Between [Vessel/Stage 1] to [Vessel/Stage 2] - At [Manufacturing Area]), [Cartridge Filter Type/Size] (Catalog No. [Filter Catalog No.]) (Between [Vessel/Stage] to [Vessel/Stage] - At [Filtration Area]), [Capsule Filter Type/Size] (Catalog No. [Filter Catalog No.]) (Between [Vessel/Stage] to [Vessel/Stage]).

- The filters are single-use filters sterilized by [Sterilization Method] (e.g., autoclaving at [Temperature Range] for [Time Duration]) as part of [Load/Process Step Name].

- Based on filter compatibility validation data, it is concluded that the filter was not affected by [Product Name] during a contact period of [Contact Duration].

- [Gas Name] was used as a processing aid for purging and blanketing operations during filtration.

- The integrity (bubble point test), chemical compatibility, adsorption, and bacterial challenge validation studies with the drug product were performed at [Facility Name, Location]. The validation studies have been summarized in [Table/Reference].

- Filter integrity parameters were established based on the bubble point validation studies.

- The integrity of the filters is tested off-line both before and after filtration.

- [Water/Other Medium] was used as a wetting medium for pre- and post-filtration integrity testing.

- Performance of the sterilizing-grade filter has been demonstrated to be acceptable, as the membrane retained [Test Organism Name] at [Concentration] per cm² of effective filtration area.

- Please refer to Table 1 summarizing the Summary of Filter Integrity Test Results for Product Name.

Table 1	Summary of Filter Integrity Test Results for Product Name

| In-Process Control | Acceptance Criteria for Exhibit Batches | Frequency | Results of the ANDA Submission Batches |  |  | Commercial batch Acceptance Criteria [BMR NO#] |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  | BATCH # 1 | BATCH # 2 | BATCH # 3 |  |
| PRE-FILTER INTEGRITY TESTING |  |  |  |  |  |  |
| POST-FILTER INTEGRITY TESTING |  |  |  |  |  |  |

Table 2	Summary of Filter Validation Studies

| Sr. No. | Type of Study | Protocol No. | Report No. |
| --- | --- | --- | --- |
| 1 | Integrity study |  |  |
| 2 | Adsorption Study |  |  |
| 3 | Chemical Compatibility study, and Bacterial Challenge study |  |  |

<Please include the results from the report in a descriptive manner>

# Component Depyrogenation

- The drug product container i.e., [Container Volume and Type] are depyrogenated using [Depyrogenation Method] in a [Depyrogenation Equipment/Facility]. Please refer to Module 3.2.P.3.3 for details regarding the design space of the component depyrogenation process i.e., depyrogenation tunnel details, depyrogenation parameters, methods and controls to monitor production cycle, load pattern and requalification or revalidation.

<Please refer below for an example for your better reference>

## Depyrogenation Validation Study Summary

A summary of the studies performed to validate the component depyrogenation process is tabulated below.

Table 3	Summary: Depyrogenation Process Validation Studies

| Equipment ID # | Validation Type | Run Type | Vial Size | Run # | Date | Reference |
| --- | --- | --- | --- | --- | --- | --- |
| XX-XXX | Initial Qualification | Empty Chamber Heat distribution run* |  | 1 | MM-DD-YYYY | X/XXX/XXX/XXXX |
|  |  |  |  | 2 | MM-DD-YYYY |  |
|  |  |  |  | 3 | MM-DD-YYYY |  |
|  |  | Heat Penetration and Endotoxin Challenge Study |  | 1 | MM-DD-YYYY | X/XXX/XXX/XXXX |
|  |  |  |  | 2 | MM-DD-YYYY |  |
|  |  |  |  | 3 | MM-DD-YYYY |  |
|  | Re-Qualification | Heat Penetration and Endotoxin Challenge Study |  | 1 | MM-DD-YYYY | X/XXX/XXX/XXXX |
|  |  |  |  | 2 | MM-DD-YYYY |  |
|  |  |  |  | 3 | MM-DD-YYYY |  |

*As per initial qualification protocol, Empty Chamber Heat distribution study is a one-time study and shall not part of requalification.

<Note: Please add the run details according to the company- and equipment-specific report. The information should match the report. The report should be provided in this section with legible graphs.>

## Validation and Production Depyrogenation Parameters

A comparative summary of the depyrogenation parameters between the validation studies and production is tabulated below.

Table 4	Depyrogenation Parameters: Validation Vs. Production for XX ml_XX mm Clear Vial

| Depyrogenation Parameters | Initial Qualification | Re-qualification | Production XX ml_XXmm Clear Vial |
| --- | --- | --- | --- |
|  | X/XXX/XXX/XXXX | X/XXX/XXX/XXXX |  |
|  | XX ml_XX mm Clear Vial | XX ml_XX mm Clear Vial |  |
| 1# heating temp. (°C) |  |  |  |
| 2# heating temp. (°C) |  |  |  |
| 3# heating temp. (°C) |  |  |  |
| 4# heating temp. (°C) |  |  |  |
| Belt Speed (mm/min.) |  |  |  |
| Air shutter at heating zone inlet open height (mm) |  |  |  |
| Air shutter at heating zone outlet open height (mm) |  |  |  |
| Cooling Temp. |  |  |  |
| Pre-heating Temp. |  |  |  |

## Initial Qualification

The XX ml_XX mm Clear Vial were initially qualified for a continuous operation load.

As part of initial qualification, three empty chamber heat distribution runs, and three heat penetration/endotoxin challenge runs were performed.

The XX ml_XX mm Clear Vial used for the proposed product Product Name.

- <Please add the container-, company-, and equipment-specific details in the write-up. All details provided here are for example purposes.>

Table 5	Vial specification: Validation Vs. Production for XX ml_XX mm Clear Vial

| - Vial Physical Dimension | Initial Qualification | Re-qualification | Product Name |  |
| --- | --- | --- | --- | --- |
|  | X/XXX/XXX/XXXX | X/XXX/XXX/XXXX | ANDA Exhibit Batch | Proposed |
|  | Vials XX mL (XX mm Collar) Clear Glass USP/In-House [Material Code: XXXXXX] |  | Vials XX mL (XX mm Collar) Clear Glass USP/In-House |  |
|  |  |  | [Material Code: XXXXXX] |  |
| Body Diameter |  |  |  |  |
| Collar outer diameter |  |  |  |  |
| Neck Diameter |  |  |  |  |
| Collar inner diameter |  |  |  |  |
| Total Height |  |  |  |  |
| Base to shoulder height |  |  |  |  |
| Neck height |  |  |  |  |
| Body wall Thickness |  |  |  |  |
| Bottom wall thickness |  |  |  |  |
| Lip thickness |  |  |  |  |
| Average overflow volume of vials |  |  |  |  |

### Load Pattern, Probes and Endotoxin Indicator Location Details

<These details are provided for reference purposes only>

<You can update this text based on the validation data>

During heat distribution runs, minimum XX number of temperature sensors equidistant throughout the width of the conveyor without touching any surfaces.

A diagrammatic presentation of the probe locations during heat distribution run is presented below:

Figure 1.	Probe Location for Heat Distribution Run

<please include image>

During heat penetration/endotoxin challenge runs, XX nests of specified containers are prepared, each nest contains one vial having sensors and one endotoxin spiked other empty vials for supporting the sensor inserted vial.

Place the temperature sensor containing vial and endotoxin spiked vials nests equidistant to cover the tunnel conveyor width at the START of the run ensuring no vials in front of vials containing temperature probes.

Initial, middle and end of run ensure the differential pressure of each chamber and chamber temperature of tunnel sterilizer is within the acceptance limit. Allow the nest vials containing temperature sensors to travel along the sterilization chamber along with endotoxin spiked vials and other vials behind.

A diagrammatic presentation of the probe location during heat penetration/endotoxin challenge runs is presented below.

Figure 2.	Probe & Endotoxin Location for Heat Penetration and Endotoxin Challenge Study: XX ml_XX mm Clear Tubular Vial

< please include image >

### Validation Study Data

The validation study results for the heat distribution and heat penetration/endotoxin challenge runs are tabulated below.

Table 6	Empty Chamber Heat Distribution Study Data

| Parameter | Report No.: X/XXX/XXX/XXXX |  |  |
| --- | --- | --- | --- |
|  | Run-1 | Run-2 | Run-3 |
|  | MM-DD-YYYY | MM-DD-YYYY | MM-DD-YYYY |
| Minimum Lethality (FH Value) |  |  |  |
| Maximum Lethality (FH Value) |  |  |  |
| Minimum Temperature in Cooling Zone (°C) |  |  |  |
| Maximum Temperature in Cooling Zone (°C) |  |  |  |
| Maximum Temp. Difference among the probes (℃) |  |  |  |

Table 7	Heat Penetration/Endotoxin Challenge Study Data: XX ml_XX mm Clear Vial

| Parameter |  | Report No.: X/XXX/XXX/XXXX |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Heat Penetration/Endotoxin Challenge |  |  |  |  |  |  |  |  |
|  |  | Run-1 |  |  | Run-2 |  |  | Run-3 |  |  |
|  |  | MM-DD-YYYY |  |  | MM-DD-YYYY |  |  | MM-DD-YYYY |  |  |
| Belt Speed |  |  |  |  |  |  |  |  |  |  |
| Minimum FH Value (minutes) |  |  |  |  |  |  |  |  |  |  |
| Maximum FH Value (minutes) |  |  |  |  |  |  |  |  |  |  |
| Minimum Temperature in Cooling Zone (°C) |  |  |  |  |  |  |  |  |  |  |
| Maximum Temperature in Cooling Zone (°C) |  |  |  |  |  |  |  |  |  |  |
| Minimum Depyrogenation time (Minutes)/probe no. |  |  |  |  |  |  |  |  |  |  |
| Maximum Depyrogenation time (Minutes)/probe no. |  |  |  |  |  |  |  |  |  |  |
| Endotoxin Results |  |  |  |  |  |  |  |  |  |  |
| Differential Pressure (Pascal) |  |  |  |  |  |  |  |  |  |  |
| Parameter |  | Run-1 |  |  | Run-2 |  |  | Run-3 |  |  |
|  |  | Initial | Middle | End | Initial | Middle | End | Initial | Middle | End |
| Pre-Heating zone |  |  |  |  |  |  |  |  |  |  |
| Heating Zone | 1st Filter |  |  |  |  |  |  |  |  |  |
|  | 2nd Filter |  |  |  |  |  |  |  |  |  |
|  | 3rd Filter |  |  |  |  |  |  |  |  |  |
|  | 4th Filter |  |  |  |  |  |  |  |  |  |
| Cooling Zone | 1st Filter |  |  |  |  |  |  |  |  |  |
|  | 2nd Filter |  |  |  |  |  |  |  |  |  |
|  | 3rd Filter |  |  |  |  |  |  |  |  |  |
| D/P between pre-heating and washing room |  |  |  |  |  |  |  |  |  |  |
| D/P between heating zone and washing room |  |  |  |  |  |  |  |  |  |  |
| D/P between cooling zone and washing room |  |  |  |  |  |  |  |  |  |  |
| D/P between washing room and filling room |  |  |  |  |  |  |  |  |  |  |

Table 8	Endotoxin Challenge Data

| INFORMATION |  |  | Report No.: X/XXX/XXX/XXXX |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | XX ml_XX mm Clear Vial |  |  |  |  |
| Manufacturer |  |  |  |  |  |  |  |
| Source |  |  |  |  |  |  |  |
| Lot No (Endotoxin Indicator Vial) |  |  |  |  |  |  |  |
| Expiry date |  |  |  |  |  |  |  |
| LRW lot no. & Expiry date |  |  |  |  |  |  |  |
| LAL (Limulus Amebocyte Lysate) Lot No. & Expiry date |  |  |  |  |  |  |  |
| Vial Size | Run # | Endotoxin Challenge (EU/Vial) | Sensitivity of LAL Reagent (EU/mL) |  | Log recovered from heat treated XX vials (EU/mL) | Log recovered from positive endotoxin indicator controls (EU/mL) | Log Reduction |
|  |  |  | Label Claim | Observed |  |  |  |
| XX ml_XX mm Clear Vial | 1 |  |  |  |  |  |  |
|  | 2 |  |  |  |  |  |  |
|  | 3 |  |  |  |  |  |  |

## Requalification

- Periodic requalification is performed on the XX ml_XX mm Clear Vial at frequency of XX year ± XX month with one run for “Heat Penetration/Endotoxin Challenge”.

### Load Pattern, Probes and Endotoxin Indicator Location Details

During heat penetration/endotoxin challenge runs, XX nests of specified containers are prepared, each nest contains one vial having sensors and one endotoxin spiked other empty vials for supporting the sensor inserted vial.

Place the temperature sensor containing vial and endotoxin spiked vials nests equidistant to cover the tunnel conveyor width at the START of the run ensuring no vials in front of vials containing temperature probes.

Initial, middle and end of run ensure the differential pressure of each chamber and chamber temperature of tunnel sterilizer is within the acceptance limit. Allow the nest vials containing temperature sensors to travel along the sterilization chamber along with endotoxin spiked vials and other vials behind.

A diagrammatic presentation of the probe location during heat penetration recipe setting and heat penetration/endotoxin challenge runs is presented below.

Figure 3.	Probe & Endotoxin Location for Heat Penetration and Endotoxin Challenge Study: XX ml_XX mm Clear Tubular Vial

<please include image>

### Validation Study Data

The validation study results for the heat penetration/endotoxin challenge runs are tabulated below.

Table 9	Heat Penetration/Endotoxin Challenge Study Data: XX ml_XX mm Clear Vial

| Parameter |  | X/XXX/XXX/XXXX |  |  |
| --- | --- | --- | --- | --- |
|  |  | Heat Penetration/Endotoxin Challenge |  |  |
|  |  | MM-DD-YYYY |  |  |
|  |  | Initial | Middle | End |
| Belt Speed |  |  |  |  |
| Minimum FH Value (minutes) |  |  |  |  |
| Maximum FH Value (minutes) |  |  |  |  |
| Minimum temperature in Cooling Zone (ºC) |  |  |  |  |
| Maximum temperature in Cooling Zone (ºC) |  |  |  |  |
| Minimum Depyrogenation time (Minutes)/probe no. |  |  |  |  |
| Maximum Depyrogenation time (Minutes)/probe no. |  |  |  |  |
| Differential Pressure (Pascal) |  |  |  |  |
| Preheating Zone |  |  |  |  |
| Heating Zone | 1st Filter |  |  |  |
|  | 2nd Filter |  |  |  |
|  | 3rd Filter |  |  |  |
|  | 4th Filter |  |  |  |
| Cooling Zone | 1st Filter |  |  |  |
|  | 2nd Filter |  |  |  |
|  | 3rd Filter |  |  |  |
| D/P between pre-heating and washing room |  |  |  |  |
| D/P between heating zone and washing room |  |  |  |  |
| D/P between cooling zone and washing room |  |  |  |  |
| D/P between washing room and filling room |  |  |  |  |

Table 10	Endotoxin Challenge Data

| INFORMATION |  |  | X/XXX/XXX/XXXX |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | XX ml_XX mm Clear Tubular Vial |  |  |  |  |
| Manufacturer |  |  |  |  |  |  |  |
| Source |  |  |  |  |  |  |  |
| Lot No (Endotoxin Indicator Vial) |  |  |  |  |  |  |  |
| Expiry date |  |  |  |  |  |  |  |
| LRW lot no. & Expiry date |  |  |  |  |  |  |  |
| LAL (Limulus Amebocyte Lysate) Lot No. & Expiry date |  |  |  |  |  |  |  |
| Vial Size | Run # | Endotoxin Challenge (EU/Vial) | Sensitivity of LAL Reagent (EU/mL) |  | Log recovered from heat treated XX vials (EU/mL) | Log recovered from positive endotoxin indicator controls (EU/mL) | Log Reduction |
|  |  |  | Label Claim | Observed |  |  |  |
| XX ml_XX mm Clear Vial | 1 |  |  |  |  |  |  |

## Acceptance Criteria

Heat Distribution

Temperature uniformity has to be ≤ [Temperature Value] in [Sterilization Chamber/ Equipment].

The FH value should be more than [Time Duration].

Heat Penetration and Endotoxin Challenge

The FH value should be more than [Time Duration].

      - A minimum endotoxin reduction of [Log Reduction Value] should be achieved at the coolest point.

      - Temperature in [Cooling Chamber/Equipment] should not be more than [Temperature Value].

      - Differential pressure of [Preheating Zone] should be [Pressure Range] Pa.

      - Differential pressure of [Heating Zone] should be [Pressure Range] Pa.

      - Differential pressure of [Cooling Zone] should be [Pressure Range] Pa.

      - Differential pressure between [Preheating Zone] and [Adjacent Room/Area] should be NMT [Pressure Value] Pa [for Requalification].

      - Differential pressure between [Heating Zone] and [Adjacent Room/Area] should be NMT [Pressure Value] Pa [for Requalification].

      - Differential pressure between [Cooling Zone] and [Adjacent Room/Area] should be NMT [Pressure Value] Pa [for Requalification].

      - Differential pressure between [Washing Room] and [Filling Room] should be [Pressure Range] Pa [for Requalification].

## Conclusion

All the Initial qualification and Requalification validation runs in the depyrogenation tunnel were meeting the acceptance criteria.

Following the successful execution of the validation runs, the operating parameters of depyrogenation tunnel for XX ml, XX mm Clear Vial were implemented for use in the manufacturing of Product Name.

# Component Sterilization

- The drug product rubber stoppers and aluminum seals are sterilized using moist heat in an autoclave.

- Please refer to Module 3.2.P.3.3 for details regarding the design space of the component sterilization process i.e. autoclave details, autoclave parameters, methods and controls to monitor production cycle, load pattern and requalification or revalidation.

## Component Sterilization Validation Study Summary

- A summary of the studies performed to validate the component sterilization process is tabulated on the following page:

Table 11	Summary: Component Sterilization Validation Studies <As example XX mm rubber stopper & XX mm seals are mentioned>

| Equipment ID# | Validation Type | Run Type | Load Description | Run# | Cycle# | Date | Reference |
| --- | --- | --- | --- | --- | --- | --- | --- |
| ABC | Initial Qualification | Empty Chamber Heat Distribution | NA | 1 |  | MM-DD-YYYY | Report Number |
|  |  |  |  | 2 |  | MM-DD-YYYY |  |
|  |  |  |  | 4 |  | MM-DD-YYYY |  |
|  |  | Heat distribution/ Heat Penetration | XX mm rubber stopper and XX mm seals (Maximum Load) | 1 |  | MM-DD-YYYY | Report Number |
|  |  |  |  | 2 |  | MM-DD-YYYY |  |
|  |  |  |  | 3 |  | MM-DD-YYYY |  |
|  |  |  | XX mm rubber stopper and XX mm seals (Minimum Load-XX) | 1 |  | MM-DD-YYYY |  |
|  |  |  |  | 2 |  | MM-DD-YYYY |  |
|  |  |  |  | 3 |  | MM-DD-YYYY |  |
|  | Re-qualification | Empty Chamber Heat Distribution | NA | 1 |  | MM-DD-YYYY | Report Number |
|  |  | Heat distribution/ Heat Penetration | XX mm rubber stopper and XX mm seals (Maximum Load) | 1 |  | MM-DD-YYYY |  |

## Validation and Production Component Sterilization Parameters

- A comparative summary of the component sterilization autoclave cycle parameters between the validation studies and production is tabulated below:

Table 12	Autoclave Cycle Parameters: Validation Vs. Production

| Autoclave Cycle Parameters | XX mm Stoppers and XX mm Seals (Program # XX) |  |  |
| --- | --- | --- | --- |
|  | Initial Qualification | Requalification | Production |
|  | Report Number | Report Number |  |
| Final Preheating Temperature (°C) |  |  |  |
| Preheating Time (minute) |  |  |  |
| Pressure (bar) |  |  |  |
| Stabilization Time (minute) |  |  |  |
| Steam Pulses Repet. |  |  |  |
| Upper Steam Pulse Pressure (bar) |  |  |  |
| Lower Steam Pulse Pressure (bar) |  |  |  |
| Sterilization set temperature (°C) |  |  |  |
| Sterilization Duration (minute) |  |  |  |
| Final Vacuum Pressure (bar) |  |  |  |
| Final Vacuum Pressure Duration (minutes) |  |  |  |

## Initial Qualification

- The XX mm rubber stoppers and XX mm seal were validated for both minimum and maximum loads. Please refer to Module 3.2.P.3.3 for the load descriptions. As part of the initial qualification, three empty chamber heat distribution runs and three heat distribution/heat penetration were performed.

### Load Pattern, Probes and Biological Indicator Location Details

- During heat distribution runs, [Number] pre-calibrated probes were placed at different locations to evaluate the uniformity of temperature distribution inside the empty chamber. Please refer to the diagram below for the probe locations during the heat distribution runs. For further details, please refer to [Report File Name / Reference].

- The cold spot was identified based on the lowest F₀ value. If the lowest F₀ value was observed at more than one location, the location with the lowest temperature was considered as the cold spot. On the identified cold spot, one external penetration probe was placed in the loaded chamber during the heat distribution/heat penetration cycles.

Figure 4.	Probe Location for Empty chamber Heat distribution Run

<Please include image>

- During maximum load heat distribution/heat penetration runs (Load # [Load Number]), a minimum of [Number] pre-calibrated probes was placed to monitor heat distribution inside the chamber, and a minimum of [Number] pre-calibrated probes were placed to monitor heat penetration inside the load components during the sterilization hold period. Accordingly, [Number] biological indicators were positioned adjacent to the heat penetration probes to assess the microbiological efficacy of the cycle.

- A diagrammatic presentation of the component arrangement, along with the probe and biological indicator locations during maximum load heat distribution/heat penetration runs, is provided below. For further details, please refer to [Report File Name / Reference].

Figure 5.	XX mm Stopper and seals (Maximum load # XXX): Load Arrangement, Probe/Biological Indicator Location for Heat Distribution/Penetration Runs

<Please include image>

- During minimum load heat distribution/heat penetration runs (Load # [Load Number]), a minimum of [Number] pre-calibrated probes was placed to monitor heat distribution inside the chamber, and a minimum of [Number] pre-calibrated probes were placed to monitor heat penetration inside the load components during the sterilization hold period. Accordingly, [Number] biological indicators were positioned adjacent to the heat penetration probes to assess the microbiological efficacy of the cycle.

- A diagrammatic presentation of the component arrangement, along with the probe and biological indicator locations during minimum load heat distribution/heat penetration runs, is provided below. For further details, please refer to [Report File Name / Reference].

Figure 6.	XX mm Stopper and seals (Minimum load # XXX): Load Arrangement, Probe/Biological Indicator Location for Heat Distribution/Penetration Runs

<Please include image>

### Validation Study Data

The validation study results for the heat distribution and heat distribution/heat penetration runs are tabulated below.

Table 13	Empty Chamber Heat Distribution Study Data

| Parameter |  |  | Report Number |  |  |
| --- | --- | --- | --- | --- | --- |
|  |  |  | Run 1 | Run 2 | Run 4 |
|  |  |  | dd-mmm-yy | dd-mmm-yy | dd-mmm-yy |
| Cycle No. |  |  |  |  |  |
| Cycle start time |  |  |  |  |  |
| Sterilization start time |  |  |  |  |  |
| Sterilization end time |  |  |  |  |  |
| Cycle end time |  |  |  |  |  |
| Sterilization Hold time |  |  |  |  |  |
| Minimum F0 value |  |  |  |  |  |
| Minimum temperature | For inbuilt probe |  |  |  |  |
|  | For external probe |  |  |  |  |
| Maximum temperature | For inbuilt probe |  |  |  |  |
|  | For external probe |  |  |  |  |
| Pressure in bar |  | Min. |  |  |  |
|  |  | Max. |  |  |  |
| Average Maximum temperature of all probes |  |  |  |  |  |
| Maximum of (Maximum-Minimum) |  |  |  |  |  |
| Maximum of (Maximum -Average) |  |  |  |  |  |
| Maximum of (Average-Minimum) |  |  |  |  |  |

Table 14	XX mm Rubber Stopper and seals (Maximum Load # XXX): Heat Distribution/Heat Penetration Data

| Parameter |  | Report Number |  |  |
| --- | --- | --- | --- | --- |
|  |  | Run 1 | Run 1 | Run 1 |
|  |  | dd-mmm-yy | dd-mmm-yy | dd-mmm-yy |
| Cycle No. |  |  |  |  |
| Cycle start time |  |  |  |  |
| Sterilization start time |  |  |  |  |
| Sterilization end time |  |  |  |  |
| Cycle end time |  |  |  |  |
| Sterilization Hold Time |  |  |  |  |
| F0 value (Min) |  |  |  |  |
| Minimum temperature (ºC) | Inbuilt probe |  |  |  |
|  | External Probe |  |  |  |
| Maximum Temperature (ºC) | Inbuilt probe |  |  |  |
|  | External Probe |  |  |  |
| Pressure in bar | Minimum |  |  |  |
|  | Maximum |  |  |  |
| Average Maximum temperature of all probes (ºC) |  |  |  |  |
| Maximum of (Maximum -Minimum) (ºC) |  |  |  |  |
| Maximum of (Maximum -Average) (ºC) |  |  |  |  |
| Maximum of (Average -Minimum) (ºC) |  |  |  |  |
| BI result |  |  |  |  |
| Chemical Indicator |  |  |  |  |

Table 15	XX mm Rubber Stopper and seals (Minimum Load # XXX): Heat Distribution/Heat Penetration Data

| Parameter |  | Report Number |  |  |
| --- | --- | --- | --- | --- |
|  |  | Run 1 | Run 1 | Run 1 |
|  |  | dd-mmm-yy | dd-mmm-yy | dd-mmm-yy |
| Cycle No. |  |  |  |  |
| Cycle start time |  |  |  |  |
| Sterilization start time |  |  |  |  |
| Sterilization end time |  |  |  |  |
| Cycle end time |  |  |  |  |
| Sterilization Hold Time |  |  |  |  |
| F0 value (min) |  |  |  |  |
| Minimum temperature (ºC) | Inbuilt probe |  |  |  |
|  | External Probe |  |  |  |
| Maximum temperature (ºC) | Inbuilt probe |  |  |  |
|  | External Probe |  |  |  |
| Pressure in bar | Minimum |  |  |  |
|  | Maximum |  |  |  |
| Average Maximum temperature of all probes (ºC) |  |  |  |  |
| Maximum of (Maximum -Minimum) (ºC) |  |  |  |  |
| Maximum of (Maximum -Average) (ºC) |  |  |  |  |
| Maximum of (Average -Minimum) (ºC) |  |  |  |  |
| BI result |  |  |  |  |
| Chemical Indicator |  |  |  |  |

Table 16	XX mm Rubber Stopper and seals for liquid Product: Microbiological Efficacy by Biological Indicator Study 	Data

| Biological Indicators Information | Report Number |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | XX mm Rubber Stopper and aluminium seals |  |  |  |  |  |
|  | Minimum Load XXX |  |  | Maximum Load XXX |  |  |  |
|  | Run 1 | Run 2 | Run 3 | Run 1 | Run 2 | Run 3 |  |
| Cycle No. |  |  |  |  |  |  |  |
| Manufacturer |  |  |  |  |  |  |  |
| Lot No |  |  |  |  |  |  |  |
| Mean population labeled |  |  |  |  |  |  |  |
| Mean population Recovered (In house) |  |  |  |  |  |  |  |
| D-value (minutes) |  |  |  |  |  |  |  |
| Z- value (°C) |  |  |  |  |  |  |  |
| Expiry date |  |  |  |  |  |  |  |
| Number of Biological indicators* exposed in each run |  |  |  |  |  |  |  |
| Biological Indicator showing growth after exposure and incubation@ |  |  |  |  |  |  |  |
| Biological Indicator Positive Control |  |  |  |  |  |  |  |
| Negative Control |  |  |  |  |  |  |  |
| Resulting Spore Log Reduction (SLR) |  |  |  |  |  |  |  |

*Strip of Biological indicator used with Geobacillus Stearothermophilus (ATCC 7953)

## Requalification

A periodic requalification is performed annually with the requalification of the maximum

- load patterns for the worst case load for the selected components as per the rationale, wherein one run for “empty chamber heat distribution” and one run of “heat distribution/heat penetration with BI challenge” study is performed.

In case of introduction of any new load, new recipe parameters are worst than that of the existing validated load, three runs for “empty chamber heat distribution” to be performed, followed by three runs for “heat distribution/heat penetration with BI challenge” study is to be performed.

A periodic requalification of steam sterilizer has been performed with one run of identified worst case maximum loads i.e. XX mm Seals and rubber stopper load and one run of “empty chamber heat distribution”. XX mm rubber stoppers and seals load (Maximum) (Load No.: - XX) is included in scheduled requalification among rubber stopper and seal of XX mm (Load No.: XX) and XX mm size load (Load No.: XX and XX) considering the same recipe parameters and highest load mass. Please refer to the report <add file name>for further details.

### Load Pattern, Probes and Biological Indicator Location Details

- During heat distribution runs, XX pre-calibrated probes were placed to check the heat distribution inside the chamber during sterilization hold period. A diagrammatic presentation of the probe location during heat distribution runs is presented below. Please refer to the report <add file name> for further details.

- Please refer to the table below for symbol identification of probes and Biological indicator in schematic load pattern diagram for heat distribution/penetration run.

Figure 7.	Probe Location for Empty chamber Heat distribution Run

<Please include image>

- During maximum load heat distribution/heat penetration runs, XX pre-calibrated probes were placed to check the heat distribution inside chamber, XX pre-calibrated probe was placed to check the heat distribution at drain location and XX pre-calibrated probes were placed to check the heat penetration inside load components during sterilization hold period. Accordingly, XX biological indicators were placed adjacent to heat penetration probes and XX biological indicator was place adjacent to drain location to check the microbiological efficacy of cycle.

- A diagrammatic presentation of the component arrangement along with the probe and biological indicator location during maximum load heat distribution/heat penetration runs is presented on the following page. Please refer to the report <add file name> for further details.

Figure 8.	 XX mm Rubber stopper and Aluminum seals (Maximum Load – XX)]: Load Arrangement, Probe/Biological Indicator Location for Heat Distribution/Penetration Runs

<Please include image>

### Validation Study Data

- The validation study results for the heat distribution and heat distribution/heat penetration runs are tabulated below.

Table 17	Empty Chamber Heat Distribution Study Data

| Parameter | Report Number |
| --- | --- |
|  | Run-1 |
|  | DD-MM-YYYY |
| Cycle No. |  |
| Cycle start time |  |
| Sterilization hold start time |  |
| Sterilization hold end time |  |
| Cycle end time |  |
| Sterilization Hold Time |  |
| Minimum temperature (ºC) |  |
| Maximum Temperature (ºC) |  |
| Minimum pressure in bar |  |
| Maximum pressure in bar |  |
| Minimum F0 value of inbuilt probe (min) |  |
| Minimum of Minimum temperature (ºC) |  |
| Maximum of Maximum temperature (ºC) |  |
| Maximum of (Maximum-Minimum) (ºC) validator probes |  |
| Maximum of (Average- Minimum) (ºC) validator probes |  |
| Maximum of (Maximum-Average) (ºC) validator probes |  |
| Minimum F0 value of external probe (min) |  |
| Equilibrium time (seconds) |  |
| Chemical Indicator color change |  |

Table 18	XX mm Rubber stopper and Aluminum seals (Maximum Load – XXX)]: Heat Distribution/Heat Penetration Data

| Parameter | Report Number |
| --- | --- |
|  | Run-1 |
|  | DD-MM-YYYY |
| Cycle No. |  |
| Cycle start time |  |
| Sterilization hold start time |  |
| Sterilization hold end time |  |
| Cycle end time |  |
| Sterilization Hold Time |  |
| Minimum temperature (ºC) |  |
| Maximum Temperature (ºC) |  |
| Minimum pressure (bar) |  |
| Maximum pressure (bar) |  |
| Minimum F0 value of Inbuilt probe (min) |  |
| Minimum of Minimum temperature (ºC) |  |
| Maximum of Maximum temperature (ºC) |  |
| Maximum of (Maximum-Minimum) (ºC) |  |
| Maximum of (Average-Minimum) (ºC) |  |
| Maximum of (Maximum-Average) (ºC) |  |
| Minimum F0 value of external probe (min) |  |
| Equilibrium time (seconds) |  |
| BI result |  |
| Chemical Indicator |  |
| Dryness of load |  |

Table 19	XX mm Rubber stopper and Aluminum seals (Maximum Load – XXX)]: Microbiological Efficacy by Biological Indicator Study Data

| Biological Indicators Information | Report Number |
| --- | --- |
|  | XX mm Rubber stopper and Aluminum seals |
|  | Maximum Load – XXX |
|  | Run 1 |
| Cycle No. |  |
| Manufacturer |  |
| Lot No |  |
| Mean population labeled |  |
| Mean population Recovered (In house) |  |
| D-value (minutes) |  |
| Z- value (°C) |  |
| Expiry date |  |
| Number of Biological indicators* exposed in each run |  |
| Biological Indicator showing growth after exposure and incubation |  |
| Biological Indicator Positive Control |  |
| Negative Control |  |
| Resulting Spore Log Reduction (SLR) |  |

*Strip of Biological indicator used with Geobacillus Stearothermophilus (ATCC 7953)

## Acceptance Criteria

Heat Distribution

- In all the probes, temperature should remain within XXX°C to XXX°C for minimum 30 minutes during sterilization hold time.

- The empty chamber heat distribution temperatures uniformity during the sterilization period (dwell time) shall be within XºC of mean chamber temperature except drain probe.

- The difference between the maximum and minimum temperature (ΔT) within the heat distribution probes (except drain point) at any time point (during sterilization phase) shall not be more than XºC.

- All the probes shall attain the equilibration time within XX seconds. Equilibration Time: The period which elapses between the attainment of the sterilization temperature in the chamber (Normally the active chamber discharge i.e., the drain) and the attainment of the sterilization temperature in all other probes at different locations.

- The recorded pressure shall be within limit (X to X bar).

- Minimum F0 value should not be less than XX minutes.

- The ink pattern on the chemical indicator sheet shall change the colour as per reference colour guide provided by manufacturer.

Heat Distribution/Heat Penetration

- All the distribution probes shall attain the equilibration time within XX seconds and all penetration probes shall attain the equilibrium time within XX seconds. Equilibration Time: The period which elapses between the attainment of the sterilization temperature in the chamber (Normally the active chamber discharge i.e. the drain) and the attainment of the sterilization temperature in all parts of the load.

- The temperature uniformity during the sterilization period (dwell time) shall be within X °C of mean chamber temperature except drain point.

- The difference between the maximum and minimum temperature (ΔT) within the probes (except drain point) at any time point shall not be more than X °C (during the exposure time).

- Sterilization hold should remain at XXX°C to XXX°C for XX minutes for all probes.

- The pressure during sterilization hold period shall be within limit (X to X bars).

- Minimum F0 value should be NLT XX minutes in all sensors.

- The chemical indicator placed in the chamber shall show necessary color change, i.e. from yellow to blue or as per vendor recommendation.

- All biological indicators (BI) exposed to the sterilization process shall be negative for growth.

- Positive BI control should show growth, and negative BI control shall not show any growth.

## Conclusion

All the run data for empty chamber heat distribution studies and product contact load heat distribution/heat penetration with BI challenge were observed to be in compliance with the acceptance criteria and stand validated.

OR

# Component Sterilization (RTU)

The drug product [Rubber Stopper Size] rubber stopper is ready-to-use, and steam sterilized by [Manufacturer Name]. For the initial qualification and requalification report of [Rubber Stopper Size] mm rubber stopper, please refer to [Manufacturer Name] DMF (DMF [DMF Number]). For DMF Letter of Authorization, please refer to Module 1.4.2 - Appendix X.

The drug product [Seal Size] mm seal is ready-to-use and supplied by [Supplier Name]. The seal is gamma sterilized by [Sterilization Facility Name]. For the initial qualification and requalification report of [Seal Size] mm Seal, please refer to [Manufacturer Name] DMF (DMF [DMF Number]). For DMF Letter of Authorization, please refer to Module 1.4.2 - Appendix X.

# EQUIPMENT Sterilization (Autoclave)

The drug product contact equipment i.e. Filling and Capping Assembly Load, are sterilized using moist heat in an autoclave.

Please refer to Module 3.2.P.3.3 for details regarding the design space of the equipment sterilization process i.e. autoclave details, autoclave parameters, methods and controls to monitor production cycle, load pattern and requalification or revalidation.

- < Additionally, ensure that details of any other autoclave used are included >

## AUTOCLAVE (XX-XXX)

### Equipment Sterilization Validation Study Summary

A summary of the studies performed to validate the equipment sterilization process is tabulated on the following page:

Table 22	Summary: Equipment Sterilization Validation Studies (Autoclave XX-XXX)

<The given details are for reference>

| Equipment ID# | Validation Type | Run Type | Load Description | Run# | Cycle# | Date | Reference |
| --- | --- | --- | --- | --- | --- | --- | --- |
| XX-XXX | Initial Qualification | Empty Chamber Heat Distribution | N/A* | 1 |  |  | X/XXX/XXX/XXXX |
|  |  |  |  | 2 |  |  |  |
|  |  |  |  | 3 |  |  |  |
|  |  | Heat distribution/ Heat Penetration | Filling and Capping Assembly Load (Maximum Load) | 1 |  |  | X/XXX/XXX/XXXX |
|  |  |  |  | 2 |  |  |  |
|  |  |  |  | 3 |  |  |  |
|  |  |  | Filling and Capping Assembly Load (Minimum Load) | 1 |  |  |  |
|  |  |  |  | 2 |  |  |  |
|  |  |  |  | 3 |  |  |  |
|  | Re- Qualification | Heat distribution/ Heat Penetration | Filling and Capping Assembly Load (Maximum Load) | 1 |  |  | X/XXX/XXX/XXXX |
| *As per initial qualification protocol, Empty Chamber Heat distribution study is a one-time study and shall not part of requalification. |  |  |  |  |  |  |  |

### Validation and Production Equipment Sterilization Parameters

- A comparative summary of the Equipment sterilization autoclave cycle parameters between the validation studies and production is tabulated below:

Table 23	Autoclave Cycle Parameters: Validation Vs. Production (Autoclave XX-XXX)

| Autoclave Cycle Parameters | Filling and Capping Assembly Load & Lyo Capping Load |  |  | Production |
| --- | --- | --- | --- | --- |
|  | Initial Qualification |  | Re-Qualification |  |
|  | X/XXX/XXX/XXXX | X/XXX/XXX/XXXX | X/XXX/XXX/XXXX |  |
| No. of Pre Pulse |  |  |  |  |
| Pre Vacuum |  |  |  |  |
| Pre Vacuum Control Band |  |  |  |  |
| Pre Vacuum Hold Time |  |  |  |  |
| Pre Vacuum Pump |  |  |  |  |
| Pre Pressure |  |  |  |  |
| Pre Pressure Control Band |  |  |  |  |
| Pre Pressure Hold Time |  |  |  |  |
| Sterile Hold Temperature |  |  |  |  |
| Control Band |  |  |  |  |
| Sterile Hold Time |  |  |  |  |
| Sterile Hold Reset |  |  |  |  |
| Sterile Hold Stop Count |  |  |  |  |
| Sterile Temperature Overshoot |  |  |  |  |
| Jacket Control Temperature |  |  |  |  |
| No. Of Post Pulse |  |  |  |  |
| Post Vacuum |  |  |  |  |
| Post Vacuum Control Band |  |  |  |  |
| Post Vacuum Hold Time |  |  |  |  |
| Post Vacuum Pump |  |  |  |  |
| Post Jacket Control Temperature |  |  |  |  |
| Door Open Temperature |  |  |  |  |
| Print Interval |  |  |  |  |
| Cond Bypass off Temperature |  |  |  |  |
| Control Temp. |  |  |  |  |

### Initial Qualification

- The Filling and Capping Assembly Load were validated for both minimum and maximum loads. Please refer to module 3.2.P.3.3 for the load descriptions.

- As part of the initial qualification, three empty chamber heat distribution runs and three heat distribution/heat penetration runs were performed.

#### Load Pattern, Probes and Biological Indicator Location Details

- During heat distribution runs, XX pre-calibrated probes were placed at different locations to check the uniformity of distribution temperature inside the empty chamber.

- Please refer to the below diagrammatic presentation of the probe location during heat distribution runs. Please refer to report 32p35-pq-xxxx-xxxx-xxxa-pq-empty-cham-auto <Add File Name> for further details.

Figure 12.	Probe Location for Empty chamber Heat Distribution Run

<Please include image.>

- During Filling and Capping Assembly maximum load heat distribution/heat penetration runs, XX pre-calibrated probes were placed to check the heat distribution inside chamber, XX pre-calibrated probes were placed to check the heat penetration inside load components during sterilization hold period. Accordingly, XX biological indicators were placed to check the microbiological efficacy of cycle.

- A diagrammatic presentation of the component arrangement during Filling and Capping Assembly maximum load heat distribution/heat penetration runs is presented below. Please refer to report 32p35-pq-xxxx-xxxx-xxxb-filling-capping-auto <Add File Name> for further details.

Figure 13.	Filling and Capping Assembly Load (Maximum Load): Load Arrangement for Heat Distribution/Penetration Runs

<Please include image>

- During Filling and Capping Assembly minimum load heat distribution/heat penetration runs, XX pre-calibrated probes were placed to check the heat distribution inside chamber, XX pre-calibrated probes were placed to check the heat penetration inside load components during sterilization hold period. Accordingly, XX biological indicators were placed to check the microbiological efficacy of cycle.

- A diagrammatic presentation of the component arrangement during Filling and Capping Assembly minimum load heat distribution/heat penetration runs is presented below. Please refer report 32p35-pq-xxxx-xxxx-xxxb-filling-capping-auto <Add File Name> for further details.

Figure 14.	Filling and Capping Assembly Load (Minimum Load): Load Arrangement for Heat Distribution/Penetration Runs

<Please include image>

#### Validation Study Data

- The validation study results for the heat distribution and heat distribution/heat penetration runs are tabulated below.

Table 24	Empty Chamber Heat Distribution Study Data (Autoclave XX-XXX)

| Parameter |  | X/XXX/XXX/XXXX |  |  |
| --- | --- | --- | --- | --- |
|  |  | Run 1 | Run 2 | Run 3 |
|  |  | MM-DD-YYYY | MM-DD-YYYY | MM-DD-YYYY |
| Cycle Number |  |  |  |  |
| Cycle start time |  |  |  |  |
| Sterilization hold start time |  |  |  |  |
| Sterilization hold end time |  |  |  |  |
| Sterilization Time (minutes) |  |  |  |  |
| Cycle end time |  |  |  |  |
| Minimum temperature observed during sterilization hold time (bar) | For Inbuilt Probe |  |  |  |
|  | For External Probe |  |  |  |
| Maximum temperature observed during sterilization hold time (bar) | For Inbuilt Probe |  |  |  |
|  | For External Probe |  |  |  |
| Minimum pressure observed during sterilization hold time (bar) |  |  |  |  |
| Maximum pressure observed during sterilization hold time (bar) |  |  |  |  |
| Maximum temperature difference from probe to probe at particular time (°C) |  |  |  |  |
| Maximum temperature difference within the probe (°C) |  |  |  |  |
| Equilibration time (seconds) |  |  |  |  |
| Minimum F0 value (minutes) |  |  |  |  |
| Cold Spot Location |  |  |  |  |
| Maximum F0 value (minutes) |  |  |  |  |
| Hot Spot Location |  |  |  |  |
| Dryness |  |  |  |  |
| Chemical Indicator |  |  |  |  |

Table 25	Filling and Capping Assembly (Maximum Load): Heat Distribution/Heat Penetration Studies Data (Autoclave XX-XXX)

| Parameter |  | X/XXX/XXX/XXXX |  |  |
| --- | --- | --- | --- | --- |
|  |  | Run 1 | Run 2 | Run 3 |
|  |  | MM-DD-YYYY | MM-DD-YYYY | MM-DD-YYYY |
| Cycle Number |  |  |  |  |
| Cycle start time |  |  |  |  |
| Sterilization hold start time |  |  |  |  |
| Sterilization hold end time |  |  |  |  |
| Sterilization Time (minutes) |  |  |  |  |
| Cycle end time |  |  |  |  |
| Minimum temperature observed during sterilization hold time (°C) | For Inbuilt Probe |  |  |  |
|  | For External Probe |  |  |  |
| Maximum temperature observed during sterilization hold time (°C) | For Inbuilt Probe |  |  |  |
|  | For External Probe |  |  |  |
| Minimum pressure observed during sterilization hold time (bar) |  |  |  |  |
| Maximum pressure observed during sterilization hold time (bar) |  |  |  |  |
| Maximum temperature difference from probe to probe at particular time (°C) |  |  |  |  |
| Maximum temperature difference within the probe (°C) |  |  |  |  |
| Equilibration time (seconds) |  |  |  |  |
| Minimum F0 value (minutes) |  |  |  |  |
| Cold Spot Location |  |  |  |  |
| Maximum F0 value (minutes) |  |  |  |  |
| Hot Spot Location |  |  |  |  |
| Chemical Indicators |  |  |  |  |
| Dryness |  |  |  |  |
| Biological Indicators (Exposed and Negative) |  |  |  |  |
| Biological indicator (positive) |  |  |  |  |

Table 26	Filling and Capping Assembly (Minimum Load): Heat Distribution/Heat Penetration Studies Data (Autoclave XX-XXX)

| Parameter |  | X/XXX/XXX/XXXX |  |  |
| --- | --- | --- | --- | --- |
|  |  | Run 1 | Run 2 | Run 4 |
|  |  | MM-DD-YYYY | MM-DD-YYYY | MM-DD-YYYY |
| Cycle Number |  |  |  |  |
| Cycle start time (HH:MM) |  |  |  |  |
| Sterilization hold start time (HH:MM) |  |  |  |  |
| Sterilization hold end time (HH:MM) |  |  |  |  |
| Sterilization Time (minutes) |  |  |  |  |
| Cycle end time (HH:MM) |  |  |  |  |
| Minimum temperature observed during sterilization hold time (°C) | For Inbuilt Probe |  |  |  |
|  | For External Probe |  |  |  |
| Maximum temperature observed during sterilization hold time (°C) | For Inbuilt Probe |  |  |  |
|  | For External Probe |  |  |  |
| Minimum pressure observed during sterilization hold time (bar) |  |  |  |  |
| Maximum pressure observed during sterilization hold time (bar) |  |  |  |  |
| Maximum temperature difference from probe to probe at particular time (°C) |  |  |  |  |
| Maximum temperature difference within the probe (°C) |  |  |  |  |
| Equilibration time (seconds) |  |  |  |  |
| Minimum F0 value (minutes) |  |  |  |  |
| Cold Spot Location |  |  |  |  |
| Maximum F0 value (minutes) |  |  |  |  |
| Hot Spot Location |  |  |  |  |
| Chemical Indicators |  |  |  |  |
| Dryness |  |  |  |  |
| Biological Indicators (Exposed and Negative) |  |  |  |  |
| Biological indicator (positive) |  |  |  |  |

Table 27	Biological Indicator Study Summary Data: Filling and Capping Assembly Load (Autoclave XX-XXX)

| Biological Indicators Information | X/XXX/XXX/XXXX |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Filling and Capping Assembly |  |  |  |  |  |
|  | Maximum Load |  |  | Minimum Load |  |  |
|  | Run 1 | Run 2 | Run 3 | Run 1 | Run 2 | Run 4 |
| Cycle No. |  |  |  |  |  |  |
| Manufacturer |  |  |  |  |  |  |
| Lot No |  |  |  |  |  |  |
| Labeled Population |  |  |  |  |  |  |
| Population Recovered (In house) |  |  |  |  |  |  |
| D-value (minutes) |  |  |  |  |  |  |
| Z- value (°C) |  |  |  |  |  |  |
| Expiry date |  |  |  |  |  |  |
| Number of Biological indicators* exposed in each run |  |  |  |  |  |  |
| Biological Indicator showing growth after exposure and incubation |  |  |  |  |  |  |
| Biological Indicator Positive Control |  |  |  |  |  |  |
| Negative Control |  |  |  |  |  |  |
| Resulting Spore Log Reduction (SLR) |  |  |  |  |  |  |

*Strip of Biological indicator used with Geobacillus Stearothermophilus (ATCC 7953)

### Requalification

Periodic requalification of the steam sterilizer is performed annually (every X year ± X month). As part of the requalification, one run each of the maximum and minimum loads (i.e., Filling and Capping Assembly Load) are conducted to perform the "Loaded Chamber Heat Distribution and Penetration Study."

#### Load Pattern, Probes and Biological Indicator Location Details

- During Filling and Capping Assembly maximum load heat distribution/heat penetration runs, XX pre-calibrated probes were placed to check the heat distribution inside chamber, XX pre-calibrated probes were placed to check the heat penetration inside load components during sterilization hold period. Accordingly, XX biological indicators were placed to check the microbiological efficacy of cycle.

- A diagrammatic presentation of the component arrangement during Filling and Capping Assembly maximum load heat distribution/heat penetration runs is presented below. Please refer to report 32p35-rq-xxxx-xxxx-xxxx-auto <Add File Name> for further details.

Figure 15.	Filling and Capping Assembly Load (Maximum Load): Load Arrangement for Heat Distribution/Penetration Runs

<Please include image>

#### Validation Study Data

- The validation study results for the heat distribution and heat distribution/heat penetration runs are tabulated below.

Table 28	Filling and Capping Assembly (Maximum Load): Heat Distribution/Heat Penetration Studies Data (Autoclave XX-XXX)

| Parameter |  | X/XXX/XXX/XXXX |
| --- | --- | --- |
|  |  | MM-DD-YYYY |
| Cycle Number |  |  |
| Cycle start time |  |  |
| Sterilization hold start time |  |  |
| Sterilization hold end time |  |  |
| Sterilization Time (minutes) |  |  |
| Cycle end time |  |  |
| Minimum temperature observed during sterilization hold time (°C) | For inbuilt sensors |  |
|  | For external probes |  |
| Maximum temperature observed during sterilization hold time (°C) | For inbuilt sensors |  |
|  | For external probes |  |
| Minimum pressure observed during sterilization hold time (bar) |  |  |
| Maximum pressure observed during sterilization hold time (bar) |  |  |
| Maximum temperature difference from probe to probe at particular time (°C) |  |  |
| Maximum temperature difference within the probe (°C) |  |  |
| Equilibration time (seconds) |  |  |
| Minimum F0 value (minutes) |  |  |
| Cold Spot Location |  |  |
| Maximum F0 value (minutes) |  |  |
| Hot Spot Location |  |  |
| Dryness |  |  |
| Chemical Indicator |  |  |
| Biological Indicator (Exposed and Negative) |  |  |
| Biological Indicators (Positive) |  |  |
| Biological Indicator Lot No. |  |  |

Table 29	Biological Indicator Study Summary Data: Filling and Capping Assembly Load (Autoclave XX-XXX)

| Biological Indicators Information | X/XXX/XXX/XXXX |
| --- | --- |
|  | Filling and Capping Assembly |
|  | Maximum Load |
| Cycle No. |  |
| Manufacturer |  |
| Lot No |  |
| Type of Biological indicator |  |
| Labeled Population CFU |  |
| Population Recovered (In house) CFU |  |
| D-value (minutes) |  |
| Z- value (°C) |  |
| Expiry date |  |
| Number of Biological indicators exposed in each run |  |
| Biological Indicator showing growth after exposure and incubation |  |
| Biological Indicator Positive Control |  |
| Negative Control |  |
| Resulting Spore Log Reduction (SLR) |  |

*Strip of Biological indicator used with Geobacillus Stearothermophilus (ATCC 7953)

## Acceptance Criteria (XX-XXX)

Heat Distribution

- There should be uniform distribution of heat in the sterilizer chamber during the sterilization hold period of XX minutes and the temperature at each temperature mapping thermocouple should be within the range of 121 °C to 124 °C during the sterilization hold period with steam pressure of XX to XX bar.

- Temperature uniformity at a given time of temperature recording between all thermocouples during hold period should not be more than X°C (temperature difference from thermocouple to thermocouple at particular time →).

- Temperature spread during hold period at a certain monitoring location should not be more than X °C (temperature difference within the thermocouple ↓).

- The equilibration time (the time difference between the first sensors reaches sterilization temperature and all sensor reaches) should not be more than XX seconds.

- The F0 Value for all the probes should be more than XX minutes.

- After completion of cycle the components should be completely dry.

- The chemical indicators should change the color as per manufacturing recommendations.

Heat Distribution/Heat Penetration

- There should be uniform penetration of heat in the load subjected for sterilization during the sterilization hold period of [Duration] minutes and the temperature at each temperature mapping probe should be within the range of [Lower Temp] to [Upper Temp] °C during the complete sterilization hold period with steam pressure of [Steam Pressure Range] bar. (Absolute Pressure: [Abs Pressure Range] bar).

- Temperature uniformity at a given time of temperature recording between all probes during hold period should not be more than [X] °C (temperature difference from probe to probe at particular time →).

- Temperature spread during hold period at a certain monitoring location should not be more than [X] °C (temperature difference within the probe ↓).

- The equilibration time (the time difference between the first sensor reach sterilization temperature and all sensor reaches) should not be more than [X] seconds.

- The F0 Value for all the probes should be more than [X] minutes.

- The chemical indicators should change the color as per manufacturing recommendation.

- No bacterial growth should be observed after the incubation period for exposed indicators and negative control.

- Growth should be observed during the incubation period for the positive control.

- After completion of cycle the components should be visually dry.

## Conclusion

- All the run data for empty chamber heat distribution studies and product contact load heat distribution/heat penetration with BI challenge were observed to be in compliance with the acceptance criteria and stand validated.

# Filtration vessel Sterilization (Sterilization in Place)

The drug product contact equipment i.e., filtration vessel is sterilized in place using moist heat.

Please refer Module 3.2.P.3.3 for details regarding the design space of the equipment sterilization in place process i.e., filtration vessel details, sterilization in place parameters, methods and controls to monitor production cycle and requalification or revalidation.

## Equipment Sterilization in Place Validation Study Summary

A summary of the studies performed to validate the equipment sterilization in place process is tabulated on the following page:

Table 30	Summary: Equipment Sterilization in Place Validation Studies

| Validation Type | Filtration vessel detail |  | Run# | Date | Reference document# |
| --- | --- | --- | --- | --- | --- |
|  | Equipment ID# | Load Description |  |  |  |
| Initial Qualification |  |  | 1 |  | X/XXX/XXX/XXXX |
|  |  |  | 2 |  |  |
|  |  |  | 3 |  |  |
| Re-qualification |  |  | 1 |  | X/XXX/XXX/XXXX |

## Validation and Production Equipment Sterilization Parameters

- A comparative summary of the equipment sterilization in place cycle parameters between the validation studies and production is tabulated below:

Table 31	Sterilization in Place Cycle Parameters: Validation Vs. Production

| Parameter | Validation Parameter |  | Production Parameter |
| --- | --- | --- | --- |
|  | Initial Qualification | Requalification |  |
|  | X/XXX/XXX/XXXX | X/XXX/XXX/XXXX | Vial Filtration Vessel |
|  | Vial Filtration Vessel - XXX L |  | XXX L |
| Heating Phase-1 Temperature (°C) |  |  |  |
| Heating Phase-1 Pressure (Bar) |  |  |  |
| Heating Phase-1 Hold Timer (Min) |  |  |  |
| Heating Phase-2 Temperature (°C) |  |  |  |
| Heating Phase-2 Pressure (Bar) |  |  |  |
| Heating Phase-2 Hold Timer (Min) |  |  |  |
| Heating Phase-3 Temperature (°C) |  |  |  |
| Heating Phase-3 Pressure (Bar) |  |  |  |
| Heating Phase-3 Hold Timer (Min) |  |  |  |
| Air Drying Cycle Start Pressure (Bar) |  |  |  |
| Sterilization Time (Min) |  |  |  |
| Sterilization Temperature (°C) |  |  |  |
| Sterilization Stop Time Temperature (°C) |  |  |  |
| Sterilization Reset Time Temperature (°C) |  |  |  |
| Air Drying End Temperature (°C) |  |  |  |
| Positive Pressure (Bar) |  |  |  |
| Air Drying Valve Open Pressure (Bar) |  |  |  |
| Air Drying Valve Close Pressure (Bar) |  |  |  |
| Positive Pressure Selection |  |  |  |
| Jacket Cooling Temperature (°C) |  |  |  |

## Initial Qualification

- The XXX L Vial filtration vessel were validated for sterilization in place.

- As part of the initial qualification, three Sterilization in Place (SIP) runs were performed.

### Probes and Biological Indicator Location Details

- During the SIP runs, pre-calibrated probes were placed at pre-defined locations to check the temperature. Correspondingly, the biological indicators were placed adjacent to the probes.

- A diagrammatic presentation of the probe and biological indicator location during the SIP runs is presented below. Please refer to report 32p35-pq-xxxx-xxxx-xxxxa-filtration-vessels <Add File Name> for further details.

Figure 16.	Diagrammatic representation of Probe/BI location for XXX L Vial filtration vessel

<Please include image>

### Validation Study Data

- The validation study results for the SIP validation runs are tabulated below.

Table 32	SIP Validation-Temperature Profile Data: XXX L Vial filtration vessel

| Parameter | X/XXX/XXX/XXXX |  |  |
| --- | --- | --- | --- |
|  | XXX L Vial filtration vessel |  |  |
|  | Run 1 | Run 2 | Run 3 |
|  | MM-DD-YYYY | MM-DD-YYYY | MM-DD-YYYY |
| Cycle start time (HH:MM) |  |  |  |
| Sterilization hold start time (HH:MM) |  |  |  |
| Sterilization hold end time (HH:MM) |  |  |  |
| Sterilization hold start time (HH:MM) |  |  |  |
| Sterilization hold end time (HH:MM) |  |  |  |
| Cycle end time (HH:MM) |  |  |  |
| Minimum temperature observed during sterilization hold time °C) |  |  |  |
| Maximum temperature observed during sterilization hold time °C) |  |  |  |
| Minimum F0 value (minutes) |  |  |  |
| Maximum F0 value (minutes) |  |  |  |

Table 33	Biological Indicators Challenge Data: XXX L Vial filtration vessel

| Biological Indicators Information | XXX L Vial filtration vessel |
| --- | --- |
|  | X/XXX/XXX/XXXX |
| Manufacturer |  |
| Lot No |  |
| Mean population labeled |  |
| Mean population Recovered (In house) |  |
| D-value (minutes) |  |
| Z- value (°C) |  |
| Expiry date |  |
| Number of Biological indicators exposed in each run |  |
| Biological Indicator Positive Control |  |
| Biological Indicator Negative Control |  |
| Resulting Spore Log Reduction (SLR) |  |

## Requalification

A periodic requalification is performed X Year ± X Month, wherein one “sterilization in place” run is to be performed.

### Probes and Biological Indicator Location Details

- During the SIP runs, pre-calibrated probes were placed at pre-defined locations to check the temperature. Correspondingly, the biological indicators were placed adjacent to the probes.

- A diagrammatic presentation of the probe and biological indicator location during the SIP runs is presented below. Please refer report 32p35-rq-xxxx-xxxx-xxxx-filtration-vessels <Add File Name> for further details.

Figure 17.	Diagrammatic representation of Probe/BI location for XXX L vessel

<Please include image>

### Validation Study Data

- The validation study results for the SIP validation runs are tabulated below.

Table 34	SIP Validation-Temperature Profile Data: XXX L Vial filtration vessel

| Parameter |  | X/XXX/XXX/XXXX |
| --- | --- | --- |
|  |  | XXX L filtration vessel |
|  |  | Run 1 |
| Cycle start time (HH:MM) |  |  |
| In-built Sensors | Sterilization hold start time (HH:MM) |  |
|  | Sterilization hold end time (HH:MM) |  |
| External Sensors | Sterilization hold start time (HH:MM) |  |
|  | Sterilization hold end time (HH:MM) |  |
| Cycle end time (HH:MM) |  |  |
| In-built Sensors | Minimum temperature observed during sterilization hold time °C) |  |
|  | Maximum temperature observed during sterilization hold time °C) |  |
| External Sensors | Minimum temperature observed during sterilization hold time °C) |  |
|  | Maximum temperature observed during sterilization hold time °C) |  |
|  | Minimum F0 value (minutes) |  |
|  | Maximum F0 value (minutes) |  |

Table 35	Biological Indicators Challenge Data: XXX L Vial filtration vessel

| Biological Indicators Information | XXX L filtration vessel |
| --- | --- |
|  | X/XXX/XXX/XXXX |
| Manufacturer |  |
| Lot No |  |
| Mean population labeled |  |
| Mean population Recovered (In house) |  |
| D-value (minutes) |  |
| Z- value (°C) |  |
| Expiry date |  |
| Number of Biological indicators exposed in each run |  |
| Biological Indicator Positive Control |  |
| Biological Indicator Negative Control |  |
| Resulting Spore Log Reduction (SLR) |  |

## Acceptance Criteria

- In all temperature sensors, minimum temperature should be NLT [Temperature] °C during sterilization hold period of [Duration] minutes.

- The F0 Value for all the thermocouples should be more than [X] minutes.

- After completion of cycle all thermocouples and Bl should remained at same locations.

- No bacterial growth should be observed after the incubation period for exposed biological indicators and negative control.

- Growth should be observed during the incubation period for the positive control of biological indicators.

## Conclusion

- All the validation data for the SIP runs of XXX L filtration vessels were observed to be complying with the set acceptance criteria and stand validated.

# LYOPHILIZer STERILIZATION (Sterilization in place)

The proposed drug product, Product Name, is lyophilized using lyophilizers manufactured by [Manufacturer Name].

Lyophilizers are proposed for the freeze-drying process, which are qualified and intended to be used for the lyophilization of the drug product.

| - Equipment Name | - Equipment ID | - Model No. | - Manufacturer | - Used for |
| --- | --- | --- | --- | --- |
| - Lyophilizer |  |  |  | - Freeze Drying of pharmaceutical formulation |

## Lyophilizer

### Equipment Sterilization in Place Validation Study Summary

- A summary of the studies performed to validate the equipment sterilization in place process is tabulated on the following page:

Table 36	Summary: Equipment Sterilization in Place Validation Studies

| Validation Type | Lyophilizer detail | Run# | Date | Reference document# |
| --- | --- | --- | --- | --- |
|  | Equipment ID# |  |  |  |
| Initial Qualification | XX-XXX | 1 | MM-DD-YYY | X/XXX/XXX/XXXX |
|  |  | 2 | MM-DD-YYY |  |
|  |  | 3 | MM-DD-YYY |  |
| Re-qualification |  | 1 | MM-DD-YYY | X/XXX/XXX/XXXX |

### Validation and Production Equipment Sterilization Parameters

- A comparative summary of the equipment sterilization in place cycle parameters between the validation studies and production is tabulated below:

Table 37	Sterilization in Place Cycle Parameters: Validation Vs. Production

| Parameter | Validation Parameter |  | Production Parameter |
| --- | --- | --- | --- |
|  | Initial Qualification | Requalification |  |
|  | X/XXX/XXX/XX | X/XXX/XXX/XX | Lyophilizer |
| Sterilization Temperature |  |  |  |
| Purge Number |  |  |  |
| Pre-evacuation pressure (BarA) |  |  |  |
| Filter Drying Time (min.) |  |  |  |
| Condenser Cooling Time (min.) |  |  |  |
| Jacket Cooling Time (min.) |  |  |  |
| Total time after set point reach (min.) |  |  |  |
| Sterilization Time (min.) |  |  |  |
| Control Pressure (BarA) |  |  |  |
| Steam Discharge Pressure (BarA) |  |  |  |
| System Drying Time (min.) |  |  |  |
| Shelf Cooling Time (min.) |  |  |  |
| Time to reach set point (min.) |  |  |  |
| Time below set point once reached (min.) |  |  |  |

### Initial Qualification

- The Lyophilizer was validated for sterilization in place.

- As part of the initial qualification, three Sterilization in Place (SIP) runs were performed.

#### Probes and Biological Indicator Location Details

- During the SIP runs, pre-calibrated probes were placed at pre-defined locations to check the temperature. Correspondingly, the biological indicators were placed adjacent to the probes.

- A diagrammatic presentation of the probe and biological indicator location during the SIP runs is presented below. Please refer to report 32p35-pq-xxxx-xxxx-xxxx-lyo <Add File Name> for further details.

Figure 18.	Diagrammatic representation of Probe/BI location for Lyophilizer

<Please include image>

#### Validation Study Data

- The validation study results for the SIP validation runs are tabulated below.

Table 38	SIP Validation-Temperature Profile Data: Lyophilizer

| Parameter | X/XXX/XXX/XX |  |  |
| --- | --- | --- | --- |
|  | Lyophilizer |  |  |
|  | Run 1 | Run 2 | Run 3 |
|  | MM-DD-YYY | MM-DD-YYY | MM-DD-YYY |
| Cycle start time (Hrs) |  |  |  |
| Sterilization hold start time (Hrs) |  |  |  |
| Sterilization hold end time (Hrs) |  |  |  |
| Sterilization Hold Time |  |  |  |
| Cycle end time (Hrs) |  |  |  |
| Minimum temperature observed during sterilization hold time (°C) |  |  |  |
| Maximum temperature observed during sterilization hold time (°C) |  |  |  |
| Minimum F0 value minutes) |  |  |  |
| Maximum F0 value minutes) |  |  |  |
| Thermocouple Location |  |  |  |
| Biological Indicators (Exposed and Negative) |  |  |  |
| Biological Indicators (Positive) |  |  |  |

Table 39	Biological Indicators Challenge Data: Lyophilizer

| Biological Indicators Information | Lyophilizer |  |  |
| --- | --- | --- | --- |
|  | X/XXX/XXX/XX |  |  |
|  | Run-1 | Run-2 | Run-3 |
| Manufacturer |  |  |  |
| Lot No |  |  |  |
| Type of Biological Indicator |  |  |  |
| Labeled Population |  |  |  |
| Population Recovered (In house) |  |  |  |
| D-value at 121°C (minutes) |  |  |  |
| Z - value (°C) |  |  |  |
| Expiry date |  |  |  |
| Number of Biological indicators exposed in each run |  |  |  |
| Biological Indicator Positive Control |  |  |  |
| Negative Control |  |  |  |
| Resulting Spore Log Reduction (SLR) |  |  |  |

### Requalification

A periodic requalification is performed X Year ± X Month, wherein one “sterilization in place” run is to be performed.

#### Probes and Biological Indicator Location Details

- During the SIP runs, pre-calibrated probes were placed at pre-defined locations to check the temperature. Correspondingly, the biological indicators were placed adjacent to the probes.

- A diagrammatic presentation of the probe and biological indicator location during the SIP runs is presented below. Please refer to report 32p35-rq-xxxx-xxxx-xxxx-lyo <Add File Name> for further details.

Figure 19.	Diagrammatic representation of Probe/BI location for Lyophilizer

<Please include image>

#### Validation Study Data

- The validation study results for the SIP validation runs are tabulated below.

Table 40	SIP Validation-Temperature Profile Data: Lyophilizer

| Parameter | Lyophilizer |
| --- | --- |
|  | X/XXX/XXX/XX |
|  | MM-DD-YYY |
| Cycle start time (Hrs) |  |
| Sterilization hold start time (Hrs) |  |
| Sterilization hold end time (Hrs) |  |
| Sterilization Hold Time |  |
| Cycle end time (Hrs) |  |
| Minimum temperature observed during sterilization hold time (°C) |  |
| Maximum temperature observed during sterilization hold time (°C) |  |
| Minimum F0 value (minutes) |  |
| Maximum F0 value (minutes) |  |
| Thermocouple Location |  |
| Biological Indicators (Exposed and Negative) |  |
| Biological Indicators (Positive) |  |

Table 41	Biological Indicators Challenge Data: Lyophilizer

| Biological Indicators Information | Lyophilizer |
| --- | --- |
|  | X/XXX/XXX/XX |
| Manufacturer |  |
| Lot No |  |
| Mean population labeled |  |
| Mean population Recovered (In house) |  |
| D-value (minutes) |  |
| Z- value (°C) |  |
| Expiry date |  |
| Number of Biological indicators exposed |  |
| Biological Indicator Positive Control |  |
| Biological Indicator Negative Control |  |
| Resulting Spore Log Reduction (SLR) |  |

# Media Fill (Aseptic Process Simulation)

## Introduction

The vial filling line located on the General Injectable Block / Vial Line of Company Name is an integrated line with filling/stoppering and sealing machine under aseptic area.

The aseptic activities carried out are routinely validated using media fill to ensure that all potential risks of aseptic processing are simulated and controlled.

The aseptic process simulation, also known as media fill, includes exposing a microbiological growth medium to product contact surfaces of equipment, container closure system and aseptic manipulations by operator or staff to closely simulate the same exposure that the product will undergo during routine production.

The sealed containers filled with liquid medium are then incubated to detect microbial contamination.

Results are interpreted to assess the potential for a unit of drug product to become contaminated during actual operations (e.g., start-up, component charging, aseptic connections, filtration, filling, sealing, etc.)

A media fill program will assess the contamination risk factors that occur on a production line and accurately assess the state of process control.

Media fill study will closely simulate aseptic manufacturing operations incorporating as appropriate, worst-case activities and conditions that may pose potential challenge to aseptic operations.

## Media Fill Validation Study Summary

- A summary of the media fill validation studies performed to validate the aseptic manufacturing process of Product Name on the General Injectable Block/ Vial Line is tabulated on the following page.

Table 42	Summary: Media Fill Validation Studies

| Location | Filling Machine No.# | Validation Type | Vial Size | Run# | Batch# | Batch Size (Theoretical) | Fill Volume (approx.) | Date | Reference |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| General Injectable Block/ Vial Line |  | Initial Qualification | X mL Vial | 1 |  |  |  |  | X/XXX/XXX/XX |
|  |  |  |  | 2 |  |  |  |  |  |
|  |  |  |  | 3 |  |  |  |  |  |
|  |  | Requalification | X mL Vial | 1 |  |  |  |  | X/XXX/XXX/XX |

## Validation and Exhibit and Commercial Equipment/Component Summary

- A comparative summary of the filling machine and container closure between the validation studies and Exhibit and Commercial is provided in the table below.

Table 43	Equipment & Component Details: Validation Vs. Exhibit and Commercial (XX mL Vial)

| Equipment/ Component Details | Initial Qualification | Exhibit | Commercial |
| --- | --- | --- | --- |
|  | X/XXX/XXX/XX |  |  |
| Filling Machine |  |  |  |
| Equipment Number |  |  |  |
| Make |  |  |  |
| Model |  |  |  |
| Container closure |  |  |  |
| Vial |  |  |  |
| Rubber Stopper |  |  |  |
| Seal |  |  |  |

The [Container Size] mL ([Neck/Collar Size]) clear tubular glass vials [USP/In-House; Material Codes: XXXXX] proposed for use in Product Name, are considered similar to the [Container Size] mL clear [Neck/Collar Size] tubular glass vials [Material Code: XXXXX] used during the initial media fill qualification. Please refer below table having comparative specification of physical dimensions for the vials.

Table 44	Comparative specification of physical dimensions for the vials

| Vial Physical Dimension | Initial Qualification | Product Name |  |
| --- | --- | --- | --- |
|  | X/XXX/XXX/XX | Exhibit Batches | Proposed |
|  | XX ml clear XX mm Collar Tubular Glass vial (Flat bottom) [Material Code: XXX] | Vials XX mL (XX mm Collar) Clear Tubular Glass USP/In-House [Material Code: XXX] | Vials XX mL (XX mm Collar) Clear Tubular Glass USP/In-House [Material Code: XXX] |
| Body Diameter |  |  |  |
| Collar outer diameter |  |  |  |
| Neck Diameter |  |  |  |
| Collar inner diameter |  |  |  |
| Total Height |  |  |  |
| Base to shoulder height |  |  |  |
| Neck height |  |  |  |
| Body wall Thickness |  |  |  |
| Bottom wall thickness |  |  |  |
| Lip thickness |  |  |  |
| Average overflow volume of Vials |  |  |  |
| Average weight of vials |  |  |  |

## Validation and Exhibit and Commercial Filling Summary

- A comparative summary of the filling duration and filling speeds between the validation studies and Exhibit and Commercial is provided in the tables below.

Table 45	Filling Summary: Validation Vs. Exhibit and Commercial: XX mL Vials (Initial Qualification)

| Filling Details | Initial Qualification |  |  | Product Name |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | X/XXX/XXX/XX |  |  | Exhibit Batch |  |  | Commercial |
|  | Batch # XXXXX | Batch # XXXXX | Batch # XXXXX | Batch # XXXXX | Batch # XXXXX | Batch # XXXXX |  |
| Vial Size |  |  |  |  |  |  |  |
| Filling Duration (Hrs.) |  |  |  |  |  |  |  |
| Filling Speed (VPM - Vial per minute) |  |  |  |  |  |  |  |
| Media |  |  |  |  |  |  |  |
| Incubation Conditions |  |  |  |  |  |  |  |

NMT: Not More Than, NA: Not Applicable

Table 46	Filling Summary: Validation Vs. Exhibit and Commercial: XX mL Vials (Requalification)

| Filling Details | Re-Qualification | Product Name |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  | X/XXX/XXX/XX | Exhibit Batch |  |  | Commercial |
|  | Batch # XXXXX | Batch # XXXXX | Batch # XXXXX | Batch # XXXXX |  |
| Vial Size |  |  |  |  |  |
| Filling Duration (Hrs.) |  |  |  |  |  |
| Filling Speed (VPM - Vial per minute) |  |  |  |  |  |
| Media |  |  |  |  |  |
| Incubation Conditions |  |  |  |  |  |

NMT: Not More Than, NA: Not Applicable

## Initial Qualification

For new facility or production process, minimum three consecutive successful process simulation runs shall be performed to release the facility/line for aseptic production.

The Vial Line at [Facility/Block Name] is new facility for both the [Dosage Form Type(s)], hence three consecutive successful process simulation tests shall be performed.

### Filled Unit Reconciliation Summary

- The filled unit reconciliation summary of the three media fill runs performed for X mL, XX mL and XX mL Vials size is tabulated below as part of the initial qualification.

Table 47	Filled Unit Reconciliation: X mL

| Stage | X/XXX/XXX/XX |  |  |
| --- | --- | --- | --- |
|  | Batch # XXXXX | Batch # XXXXX | Batch # XXXXX |
| Good Vials received for Incubation (Nos.) |  |  |  |
| Good vials incubated (Population-1) |  |  |  |
| Good vials incubated (Population-2) |  |  |  |
| No. of Good Vials after 7 days Incubation (Pop 1 + Pop 2) |  |  |  |
| No. of Turbid Vials after 7 days Incubation |  |  |  |
| Minimum Temperature during Incubation (20-25 °C) |  |  |  |
| Maximum Temperature during Incubation (20-25 °C) |  |  |  |
| No. of Good Vials after 14 days Incubation (Pop 1 +Pop 2) |  |  |  |
| No. of Turbid Vials after 14 days Incubation |  |  |  |
| Minimum Temperature during Incubation (30-35 °C) |  |  |  |
| Maximum Temperature during Incubation (30-35 °C) |  |  |  |
| Good Vials Accountability |  |  |  |
| Growth observed in filled units |  |  |  |
| Result of post GPT |  |  |  |
| Result of CCIT (Container Closure Integrity Test) |  |  |  |

- The environmental monitoring test results (Viable and non-viable) were observed to be complying with the acceptance criteria. Please refer to the media fill reports i.e. 32p35-media-fill-report-initial-Xml <Add File Name> for the detailed environmental monitoring data.

## Requalification

- Periodic: When there are no major changes to a validated status of aseptic manufacturing and filling with respect to area, equipment, processes and systems, a single successful media fill run shall be taken on each filling line every X Months from date of filling ± X month.

- Following shall be strategy for periodic requalification:

- Yearly- Minimum OR maximum size container alternatively (worst case containers)

- Yearly- Remaining intermediate size container at least once in a year. Where there are many containers with multiple sizes then it shall be ensured that all intermediate size containers shall be covered within 5 years and rotated on 5 yearly basis.

### Filled Unit Reconciliation Summary

- Tabulated below is the filled unit reconciliation summary of the one media fill run performed for X mL Vials size.

Table 48	Filled Unit Reconciliation: X mL

| Stage |  | X/XXX/XXX/XX |
| --- | --- | --- |
|  |  | Batch # XXXXX |
| Number of Vials received for zero days Visual Inspection (Nos) |  |  |
| Rejected Vials after zero days Visual Inspection (Nos) |  |  |
| Good vials incubated (Nos) |  |  |
| No. of Good Vials after 7 days Incubation (Nos) |  |  |
| No. of Turbid Vials after 7 days Incubation (Nos) |  |  |
| Temperature during Incubation (20-25 °C) | Minimum (°C) |  |
|  | Maximum (°C) |  |
| No. of Good Vials after 14 days Incubation (Nos) |  |  |
| No. of Turbid Vials after 14 days Incubation (Nos) |  |  |
| Temperature during Incubation (30-35 °C) | Minimum (°C) |  |
|  | Maximum (°C) |  |
| Good Vials Accountability |  |  |
| Growth observed in filled units |  |  |
| Result of post GPT |  |  |
| Result of CCIT (Container Closure Integrity Test) |  |  |

- The environmental monitoring test results (Viable and non-viable) were observed to be complying with the acceptance criteria. Please refer to the media fill reports i.e. 32p35-media-fill-report-recent-xml <Add File Name> for the detailed environmental monitoring data.

## Acceptance Criteria

- Population II: Population-II container are incubated during the media fill, for study purpose only to know the probable reasons, as in routine production these containers are discarded, hence no acceptance criteria are defined for population-II containers.

- Population I: Based on the quantity of the confirmed contaminated container, following is the acceptance criteria to determine if the media fill passes, when to conduct an investigation and repeat the media fill. as per below table.

<Below details are provided for reference purposes only>

| - No. of filled units to be incubated | - Action Item |
| --- | --- |
| - When Less than 5000 | - 1 confirmed contaminated unit is considered cause for revalidation following an instigation. |
| - When Between 5000 - 10,000 | - 1 confirmed contaminated should result in an investigation, including consideration of a repeat media fill. |
|  | - 2 confirmed contaminated unit is considered cause for revalidation following an instigation. |
| - When More than 10,000 | - 1 confirmed contaminated unit requires an investigation |
|  | - 2 confirmed contaminated unit is considered cause for revalidation following an instigation. |

- In case of any confirmed single contaminated container, an investigation shall be conducted by production and cross-functional team including Microbiologist to find out the source of contamination and the investigation report must justify why the media fill shall not be repeated.

## Conclusion

- Based on the observed study results, it was concluded that the proposed aseptic manufacturing process for XX mL vial sizes stands validated.

## Media Fill Failure Investigation and Corrective Action

<Please include Media fill failure Investigation and Correction Action from SOP>

<Please note that for terminal sterilization, the following points need to be added if terminal   sterilization method is used>

# Terminal Sterilization

<Below information about the terminal sterilization is provide for your reference.>

The drug product filled vials are sterilized using steam air mixture using F0 based (F0 value: NLT XX minutes) Terminal Moist Heat sterilization cycle in a steam air sterilizer. The sterilization of drug product filled vials is achieved through transfer of heat from steam air mixture to the container (vial) exterior and then to the aqueous liquid product contained within a steam air sterilizer.

Terminal sterilization was performed for this product and D-value calculated from a survivor curve was XX minutes at XXX°C. Terminal Sterilization suitability study at XXX°C was carried, analytical results indicated that Terminal Sterilization is suitable for the Product Name as discussed in Section 2.2.1.10 of Module 3.2.P.2. Terminal Sterilization cycle was validated to achieve F0 value of > XX minutes.

For terminal sterilization qualification studies:

- Heat distribution in steam air mixture was performed at 121°C to 124°C for minimum XX minutes.

- Loaded Chamber Heat Penetration with Biological indicator study was carried out as per approved protocol and achieved F0 value NLT XX min.

For terminal sterilization of Product Name the operating parameter used for the validation runs ensure that the load is exposed to achieve F0 NLT XX minutes.

Please refer Module 3.2.P.3.3 for details regarding the design space of the Terminal sterilization process i.e., Steam sterilizer details, Steam sterilizer parameters, methods and controls to monitor production cycle, load pattern and requalification or revalidation.

## Terminal Sterilization Validation Study Summary

A summary of the studies performed to validate the component Terminal Sterilization process is tabulated on the following page:

Table 49	Summary: Terminal Sterilization Process Validation Studies

| Equipment Id # | Validation Type | Run Type | Load Description | Run # | Date | Reference |
| --- | --- | --- | --- | --- | --- | --- |
| XXXXXXX | Initial Qualification | Empty Chamber Heat distribution study | XX | 1 | MM-DD-YYYY | Report # XXX |
|  |  |  |  | 2 | MM-DD-YYYY |  |
|  |  |  |  | 3 | MM-DD-YYYY |  |
|  |  | Heat Penetration Study with Biological Indicators challenge test | XX mL vial (Maximum load #XX) | 1 | MM-DD-YYYY | Report # XXX |
|  |  |  |  | 2 | MM-DD-YYYY |  |
|  |  |  |  | 3 | MM-DD-YYYY |  |
|  |  |  | XX mL vial (Minimum load #XX) | 1 | MM-DD-YYYY |  |
|  |  |  |  | 2 | MM-DD-YYYY |  |
|  |  |  |  | 3 | MM-DD-YYYY |  |

## Validation and Production Terminal Sterilization Parameters

A comparative summary of the Terminal Sterilization parameters between the validation studies and production is tabulated below.

Table 50	Terminal Sterilization Parameters: Validation Vs. Production

| Parameters | Set Value |  |
| --- | --- | --- |
|  | Terminal Sterilization for XX ML Vial |  |
|  | Initial Qualification | Production |
|  | Report No. | Report No |
| Pre Heat Temperature (°C) |  |  |
| Pre Heat Time (Sec) |  |  |
| Sterilize Temperature (°C) |  |  |
| Sterilize Time (Sec) |  |  |
| F0 Value (min) |  |  |
| Slow Cool Temperature (°C) |  |  |
| Fast Cool Temperature (°C) |  |  |
| Heat Pressure Coef |  |  |
| Ster Pressure Coef |  |  |
| Cool Pressure Coef |  |  |
| High Temp Alarm SV (°C) |  |  |
| Low Temp Alarm SV (°C) |  |  |
| Steam Large Close Temp (°C) |  |  |
| PID PARAMETERS |  |  |
| PID Temp 1 Set (°C) |  |  |
| PID Out 1 Set (%) |  |  |
| PID Temp 2 Set (°C) |  |  |
| PID Out 2 Set (%) |  |  |
| PID Temp 3 Set (°C) |  |  |
| PID Out 3 Set (%) |  |  |
| PID Temp 4 Set (°C) |  |  |
| PID Out 4 Set (%) |  |  |

## Initial Qualification

The product filled vials were validated for both minimum and maximum loads. Please refer to Module 3.2.P.3.3 for the load descriptions. As part of the initial qualification, three runs of empty chamber heat distribution and three runs of Heat Penetration Study with Biological Indicators challenge test were performed.

### Load Pattern, Probes and Biological Indicator Location Details

During empty chamber heat distribution runs, 18 probes were placed at different location to check the uniformity of distribution of temperature inside the empty chamber and 2 probe were placed, each near drain loading and unloading side. A diagrammatic presentation of the probe location during heat distribution runs is presented below. Please refer report 32p35-xxxx-xxxx-xxxx <Add file Name> for further details.

Figure 20.	Probe Location for Empty chamber Heat Distribution Run

<Please include image>

For Terminal Sterilization Cycle XX mL PP Vial with XX mL fill volume, during maximum loaded chamber qualification study (Load# XX) XX probes were placed (including X probes near inbuilt temperature product probes T1 and T2) to check the heat penetration inside center of filled vial during sterilization hold period. Additionally, XX biological indicators were inserted in PP vial and were placed adjacent to the PP vial containing heat penetration probe.

For Terminal Sterilization Cycle XX mL PP Vial with X mL fill volume, during minimum loaded chamber qualification study (Load# X), X probes were placed (including X probes near inbuilt temperature product probes T1 and T2) to check the heat penetration inside centre of filled vial during sterilization hold period. Additionally, X biological indicators were inserted in PP vial and were placed adjacent to the PP vial containing heat penetration probe.

Please refer report to 32p35-xxxx-xxxx-xx (Annexure-IA) <Add file Name> for further details.

### Validation Study Data

The validation study results for the heat distribution and heat penetration/biological indicator challenge runs are tabulated below.

Table 51	Empty Chamber Heat Distribution Study Data (Load # XX)

| Parameter |  | Report # |  |  |
| --- | --- | --- | --- | --- |
|  |  | Run-1 | Run-2 | Run-3 |
|  |  | MM-DD-YYYY | MM-DD-YYYY | MM-DD-YYYY |
| Cycle start time |  |  |  |  |
| Cycle end time |  |  |  |  |
| Sterilization hold start time |  |  |  |  |
| Sterilization hold end time |  |  |  |  |
| Total sterilization hold time |  |  |  |  |
| Sterilization Dwell Time (minute) |  |  |  |  |
| Minimum temperature (ºC) | For inbuilt probe |  |  |  |
|  | For external probe |  |  |  |
| Maximum temperature (ºC) | For inbuilt probe |  |  |  |
|  | For external probe |  |  |  |
| Minimum F0 value (min) | For inbuilt probe |  |  |  |
|  | For external probe |  |  |  |
| Maximum F0 value (min.) | For external probe |  |  |  |

Table 52	XX mL PP Vial (Maximum Load# XX): Heat Penetration Study with BI challenge run

| Parameter |  | Report # |  |  |
| --- | --- | --- | --- | --- |
|  |  | Run-1 | Run-2 | Run-3 |
|  |  | MM-DD-YYYY | MM-DD-YYYY | MM-DD-YYYY |
| Cycle start time |  |  |  |  |
| Cycle end time |  |  |  |  |
| Maximum Temperature (ºC) |  |  |  |  |
| Minimum F0 value (min.) | For inbuilt probe |  |  |  |
|  | For external probe |  |  |  |
| Maximum F0 value (min.) | For external probe |  |  |  |
| BI Result |  |  |  |  |
| CCIT by dye ingress method |  |  |  |  |

Table 53	XX mL Vial (Minimum Load# XX): Heat Penetration Study with BI challenge run

| Parameter |  | Report # |  |  |
| --- | --- | --- | --- | --- |
|  |  | Run-1 | Run-2 | Run-3 |
|  |  | MM-DD-YYYY | MM-DD-YYYY | MM-DD-YYYY |
| Cycle start time |  |  |  |  |
| Cycle end time |  |  |  |  |
| Maximum Temperature (ºC) |  |  |  |  |
| Minimum F0 value (min.) | For inbuilt probe |  |  |  |
|  | For external probe |  |  |  |
| Maximum F0 value (min.) | For external probe |  |  |  |
| BI Result |  |  |  |  |

Table 54	Microbiological Efficacy by Biological Indicator Study Data

| Biological Indicators Information | Strength 1 PP Vial (Report#) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Maximum Load# XX |  |  | Minimum Load# XX |  |  |
|  | Run- 1 | Run- 2 | Run- 3 | Run- 1 | Run- 2 | Run- 3 |
| Manufacturer |  |  |  |  |  |  |
| Lot No |  |  |  |  |  |  |
| Mean population labeled |  |  |  |  |  |  |
| Mean population Recovered (In house) |  |  |  |  |  |  |
| D-value (minutes) |  |  |  |  |  |  |
| Z value |  |  |  |  |  |  |
| Manufacturing date |  |  |  |  |  |  |
| Expiry date |  |  |  |  |  |  |
| Number of Biological indicators exposed in each run* |  |  |  |  |  |  |
| Biological Indicator showing growth after exposure and incubation |  |  |  |  |  |  |
| Biological Indicator Positive Control |  |  |  |  |  |  |
| Negative Control |  |  |  |  |  |  |

* Biological indicator: Geobacillus Stearothermophilus (ATCC 7953)

## Requalification

Annually one run of Empty Chamber Heat Distribution (ECHD) Study shall be carried out to evaluate inherent capability of terminal sterilization to achieve temperature stability and uniformity of the empty chamber during Sterilization hold period. In case of any change in recipes, three runs of Empty Chamber Heat Distribution (ECHD) Study shall be carried out.

Annually one run for Loaded Chamber Heat Penetration (HP) with BI Challenge study for identified loads (Maximum Load only) shall be carried out to verify that minimum lethality is achieved in the product container throughout the load inside the chamber and product contents inside containers (vials) are sterilized. In case of introduction of new load / revised load / change in recipe or as per CCF assessment, three runs shall be carried out.

## Acceptance Criteria

Heat distribution

- In all the sensor temperature should remain between 121° C to 124° C for minimum XX minute during sterilization hold time.

- F0 value should not be less than XX minutes.

- The difference between the maximum and minimum temperature (ΔT) within the heat

- The difference between the maximum and minimum temperature (ΔT) within the heat is not more than 2°C.

Heat Penetration and Endotoxin Challenge

- F0 value should not be less than XX minutes in all external and inbuilt temperature sensors.

- Cumulative maximum F0 value is for information only.

- The exposed biological indicator should not show any growth.

- Container closure integrity test should pass.

## 10.6	Conclusion

All the run data for empty chamber heat distribution studies were observed in compliance with the acceptance criteria and stand validated.

All the run data for product fill loaded heat penetration with BI challenge and Container closure integrity test were observed in compliance with the acceptance criteria and stand validated.

<Note: The points below are not applicable for terminal sterilization>

5. Equipment Sterilization (Autoclave)

6. Filtration Vessel Sterilization (Sterilization in Place)

7. Lyophilizer Sterilization (Sterilization in Place)

8 Media Fill

# Container Closure Integrity validation

## Introduction

Container Closure Integrity testing is performed to ensure container closure system defined for the product Name is validated to provide a sterile barrier to product and product safety is maintained throughout shelf life.

The container closure system used in the exhibit batch and proposed for commercial batch manufacturing is given in the table below.

Table 55	Container closure system for Product Name

| Packaging Configuration | Type of Materials | Used in Exhibit & Commercial Batch | Material Code | Description |
| --- | --- | --- | --- | --- |
| Primary Packaging Material | Vial | Used in Exhibit Batch |  |  |
|  |  | For Proposed Commercial Batch |  |  |
|  | Rubber Stopper | Used in Exhibit Batch |  |  |
|  |  | For Proposed Commercial Batch |  |  |
| Secondary Packaging Materials | Seal | Used in Exhibit Batch and For Proposed Commercial Batch |  |  |

For procedure of Container Closure Integrity test, refer Module 3.2.P.5.2 i.e. 32p52-method and for validation of Container Closure Integrity test, refer Module 3.2.P.5.3 i.e., 32p53-validation-ccit-method.

Table 56	Acceptance Criteria and observation: Dye Ingress-ANDA Submission Batch of Product Name

| Acceptance criteria | Observation |  |  |
| --- | --- | --- | --- |
|  | Batch # XXXXX | Batch # XXXXX | Batch # XXXXX |
| None of vial should fail in the leak test | None of vial is fail in the leak test. |  |  |

## Conclusion

Container closure integrity test by dye ingress test shows satisfactory results and in line with acceptance criteria. It is concluded that the proposed container closure system and sealing parameter derived after validation are suitable to ensure sterile barrier for the Product Name. Hence the proposed container closure system of the drug product stands validated to function as a barrier against microbial ingress.

## Methods and controls to monitor production cycles

- During routine production, validated parameters like machine speed is predefined in the Batch Manufacturing Record. These parameters are recorded by production. As an in-process control, Capping quality is checked at start, at interval of every XX ± X minutes and at the end of operation and record the results of Quality of Capping. After completion of capping activity, samples are collected and vacuum leak testing of collected vials at start, middle and end of Capping operation by manufacturing team. The result of Container Closure Integrity (Leak test) test is recorded in Finished Product certificate of analysis and reviewed by Quality Assurance person as a part of batch release.

# Exhibit Batch Filling Data

- [Company Name] has generated the data for the filling stage at the initial, middle, and end of filling. The tests Description, pH, Identification by HPLC, Assay (by HPLC), and Related Substances by HPLC are analyzed at each of these stages. The results are summarized in the table below.

Table 57	Exhibit Batch filling results of Product Name (Filling stage: Initial, Middle & End)

| Sr. No | In-process Control | Acceptance criteria | Batch # XXXXX |  |  | Batch # XXXXX |  |  | Batch # XXXXX |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Initial | Middle | End | Initial | Middle | End | Initial | Middle | End |
| 1 |  |  |  |  |  |  |  |  |  |  |  |
| 2 |  |  |  |  |  |  |  |  |  |  |  |
| 3 |  |  |  |  |  |  |  |  |  |  |  |

# Holding periods

Stage wise hold time studies have been performed to establish the maximum allowable hold times for the unfiltered bulk and filtered bulk. Details on the hold time study stages, storage condition, testing interval and tests performed have been tabulated below:

Table 58	Details of Holding Time Study

| Stage | Testing Interval | Test Performed |
| --- | --- | --- |

- The Hold time study results of the executed batch are tabulated on the following pages.

Table 59	Exhibit Batch bulk hold time results of Product Name (Unfiltered bulk (store at 60°C): Bulk solution 0 hour, 12 hours, 24 hours & 36 hours hold from end of compounding)

| Sr. No. | In-process Control | Acceptance criteria | Batch # XXXXX |  |  |  | Batch # XXXXX |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | 0 hrs | 12 hrs | 24 hrs | 36 hrs | 0 hrs | 12 hrs | 24 hrs | 36 hrs |
| NA – Not Applicable, NMT - Not more than; ND -Not detected; BQL- Below Quantification Limit; BDL- Below Detection Limit |  |  |  |  |  |  |  |  |  |  |

Table 60	Exhibit Batch bulk hold time results of Product Name (Filtered bulk (Store at 60°C): Bulk solution 0 hour, 12 hours, 24 hours & 36 hours hold)

| Sr. No. | In-process Control | Acceptance criteria | Batch # XXXXX |  |  |  | Batch # XXXXX |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | 0 hrs | 12 hrs | 24 hrs | 36 hrs | 0 hrs | 12 hrs | 24 hrs | 36 hrs |
| NA – Not Applicable, NMT - Not more than; ND -Not detected; BQL- Below Quantification Limit; BDL- Below Detection Limit |  |  |  |  |  |  |  |  |  |  |

Table 61	Exhibit Batch bulk hold time results of Product Name (Filtered bulk (Store at NMT 25°C): Bulk solution 0 hour, 12 hours, 24 hours & 36 hours hold)

| Sr. No. | In-process Control | Acceptance criteria | Batch # XXXXX |  |  |  | Batch # XXXXX |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | 0 hrs | 12 hrs | 24 hrs | 36 hrs | 0 hrs | 12 hrs | 24 hrs | 36 hrs |
| NA – Not Applicable, NMT - Not more than; ND -Not detected; BQL- Below Quantification Limit; BDL- Below Detection Limit |  |  |  |  |  |  |  |  |  |  |

## Conclusion

- Based on the data generated as tabulated above during hold time studies of the executed batch, stage wise hold time has been proposed. Please refer to Module 3.2.P.3.4 for stage wise hold time proposal.